## Slide 1: Development of a Silicon Drift Detector Array to Search for keV-scale Sterile Neutrinos with the KATRIN Experiment
Detector not to scale!

## Slide 2: Tritium and Neutrinos
-decay of tritium
Maximal electron energy reduced by neutrino mass

## Slide 3: Neutrino Mass Measurement
Integral measurement (MAC-E) @ endpoint
Effective neutrino mass
eV (90% CL)
KATRIN Collab. Nat. Phys. 18, 160–166 (2022)
MAC-E spectrometer
retarding
potential
e-
1 e/s

## Slide 4: Neutrino Mass Measurement
Integral measurement (MAC-E) @ endpoint
Effective neutrino mass
eV (90% CL)
MAC-E spectrometer
retarding
potential
e-
Chart: 
| m3 | m2 | m1
Mixing | 0.2 | 0.3 | 0.5
Chart: 
| m3 | m2 | m1
Mixing | 0.2 | 0.3 | 0.5
KATRIN Collab. Nat. Phys. 18, 160–166 (2022)
1 e/s

## Slide 5: What are sterile Neutrinos?
Neutrinos are only left-handed
Right-handed neutrinos are a minimal extension to SM
No weak interaction -> sterile
Additional mass eigenstates
Chart: 
| m3 | m4 | m2 | m1
Mixing | 0.2 | 0.05 | 0.3 | 0.5
Chart: 
| m3 | m4 | m2 | m1
Mixing | 0.2 | 0.05 | 0.3 | 0.5
Chart: 
| m3 | m2 | m1 | m4
Mixing | 0.1 | 0.1 | 0.1 | 2
Chart: 
| m3 | m2 | m1 | m4
Mixing | 0.1 | 0.1 | 0.1 | 2
    keV sterile -> dark matter candidate

## Slide 6: Beta Decay also Sensitive to keV Sterile Neutrinos
Tiny kink-like imprint in -decay
High rates required
Installation planned after -mass campaign completed end of 2025
Detector not to scale!
Kink-like signature
108 e/s
TRISTAN Project = high precision electron spectroscopy detector upgrade for the KATRIN experiment

## Slide 7: Beta Decay also Sensitive to keV Sterile Neutrinos
Spectral distortion in -decay
Kink-like signature
108 e/s

## Slide 8: Beta Decay also Sensitive to keV Sterile Neutrinos
Spectral distortion in -decay
High rates required

## Slide 9: Beta Decay also Sensitive to keV Sterile Neutrinos
Spectral distortion in -decay
High rates required
Installation planned after -mass campaign completed end of 2025
Detector not to scale!
108 e/s
TRISTAN SDD = high precision electron spectroscopy detector upgrade for the KATRIN experiment
Kink-like signature

## Slide 10: Challenges for TRISTAN Detector System
High rates )
  1000 pixels (100 kcps per pixel)
  Flux-tube Ø 20 cm -> Ø 3mm pixel
Differential measurement
  Energy resolution < 300 eV (@20 keV)
Low energy threshold < 2 keV
  Thin entrance window (< 100 nm)
    Silicon Drift Detector (SDD)

## Slide 11: Challenges for TRISTAN Detector System
    Silicon Drift Detector (SDD)
    9 modules with each 166 pixels
High rates )
  1000 pixels (100 kcps per pixel)
  Flux-tube Ø 20 cm -> Ø 3mm pixel
Differential measurement
  Energy resolution < 300 eV (@20 keV)
Low energy threshold < 2 keV
  Thin entrance window (< 100 nm)
40 mm

## Slide 12: Development of the first TRISTAN Detector Module
Commissioning Tests
Realistic Environment
Assembly Procedure
How to assemble the detector module?
Characterization with X-Rays in ‘perfect lab’ conditions
Testing the module in ‘Mini-KATRIN’ (MoS) with electrons

## Slide 13
Critical Assembly Step – Gluing SDD to Cesic
SDD is glued to Cesic heat exchange
  Same thermal expansion as silicon (SDD)
First attempts less than 70% yield
  Investigations to locate issue
  Swapping to glue without spacer pills
Daniel Siegmann
13
S0-166-1 (D.Va)
S0-166-2 (Mercy)
S0-166-3(Echo)
S0-166-4 (Winston)
Gluing procedure works (>99% yield) 
Notes: 13

## Slide 14: Development of the first TRISTAN Detector Module
Commissioning Tests
Realistic Environment
Assembly Procedure
How to assemble the detector module?
Characterization with X-Rays in ‘perfect lab’ conditions
Testing the module in ‘Mini-KATRIN’ (MoS) with electrons
ü

## Slide 15: Characterization TRISTAN Module (Winston)
Detector installed in test setup @ MPP
  Optimized detector settings
  Pixel 7 has a loose contact
Investigated with X-ray source (55Fe)
Energy resolution (FWHM)
X-ray @ 5.9 keV
ü

## Slide 16: Operation Temperature for Detector Module
Current operation point
Electronic noise reduces with temperature
Current operation:
  Ciller set between -50°C and -70°C
  No large improvement expected for lower temperatures
  Currently also not possible to go lower
Target temperature for detector modules roughly -30°C @ SDD

## Slide 17: Development of the first TRISTAN Detector Module
Commissioning Tests
Realistic Environment
Assembly Procedure
How to assemble the detector module?
Characterization with X-Rays in ‘perfect lab’ conditions
Testing the module in ‘Mini-KATRIN’ (MoS) with electrons
ü
ü

## Slide 18
Testing in KATRIN-like conditions
  Vacuum compatibility
  Magnetic compatibility
  Cabling and interconnections
  Understanding electronic noise
  Experience with KATRIN framework
source for commissioning
TRISTAN Module in Monitor-Spectrometer
Daniel Siegmann
18
Detector not to scale!
retarding
potential
e-
Notes: 18

## Slide 19: Commissioning with X-Ray Source
Successful installation end of 2022
Measured spectrum inside MoS
‘Detector response’
X-ray @ 5.9 keV
C. Bruch
Detector functional and comparable X-ray performance

## Slide 20
Testing in KATRIN-like conditions
  Vacuum compatibility
  Magnetic compatibility
  Cabling and interconnections
  Understanding electronic noise
  Experience with KATRIN framework
source for commissioning
Mono-energetic electron
  source: 7 keV to 32 keV
  Accelerated electrons from wall
TRISTAN Module in Monitor-Spectrometer
Daniel Siegmann
20
Detector not to scale!
retarding
potential
e-
source
Notes: 20

## Slide 21: Measurements with Electrons
source used with retarding potential
Detector compliant with MAC-E filter
Overall good performance, but …..
‘Detector response’
e-

## Slide 22: Inhomogeneity in Energy Resolution for Electrons
Electron @ 18 keV
Electron @ 31 keV
30 eV
e-
e-

## Slide 23: Inhomogeneity in Energy Resolution for Electrons
Various tests exclude MoS as cause of inhomogeneity
  Excluded environment, e.g. light, electronic noise, angle of electrons …
Electron @ 18 keV
30 eV
e-
Effect must be related to the SDD chip
Electron @ 31 keV
e-

## Slide 24: Effect of the Entrance Window
Fraction of charge/signal is lost in entrance window

## Slide 25: Effect of the Entrance Window
Fraction of charge/signal is lost in entrance window
Inhomogeneity caused by ticker entrance window
Chip homogeneity expected to be better
Sensitive area
Entrance window

## Slide 26: Development of the first TRISTAN Detector Module
Commissioning Tests
Realistic Environment
Assembly Procedure
How to assemble the detector module?
Characterization with electrons in ‘perfect lab’ conditions
Testing the module in ‘Mini-KATRIN’ (MoS) with electrons
ü
ü

## Slide 27: Cleaning the Detector Surface
Electron gun @ 5 keV
Additional cleaning of detector
  Isopropanol & Acetone
Improved detector on most regions
  Worsened energy resolution in one spot
Baseline @ 5 keV
Cleaning
Cleaned @ 5keV

## Slide 28: Cleaning the Detector Surface
Electron gun @ 5 keV
Additional cleaning of detector
  Isopropanol & Acetone
Improved detector on most regions
  Worsened energy resolution in one spot
  Corresponds to waver center
  Seen in 2 independent detectors
Baseline @ 5 keV
Cleaned @ 5keV
Substance on the surface causes inhomogeneity
Can be cleaned away 

## Slide 29: Off-Center Detector Measurements
S0-166-6 @ 5keV
S0-166-4 @ 5keV
Chips from the off-center of the waver very homogeneous
Can be used for the final detector assembly

## Slide 30: Energy resolution for X-Rays and Electrons
FWHM for X-ray follow Fano statistics
FWHM for electrons in good region
  Entrance window effects visible
  Approx. 40 to 50 eV higher than X-ray
electron slightly larger
  Additional scattering in source
  Mimics entrance window effect
Electron data only in ‘clean region’

## Slide 31: Towards the TRISTAN Detector System
TRISTAN Module
166 Pixels Detector
Commissioning successful
Fulfills requirements
TRISTAN Detector (Phase 1)
9 x 166 Pixels 1000 Pixels
Integration after completion of -mass measurement (2025)
More than scaling up the detector
Detector chamber
High density electronics
…
ü

## Slide 32
... and many more!
C. Bruch
M. Gugiatti
D. Siegmann
K. Urban
M. Carminati
T. Houdy
M. Descher
S. Lichter
P. King
J. Wolf
M. Steidl
S. Mertens
D. Hinz
F. Edzards
L. Wunderl
C. Forstner
D. Spreng
P. Lechner
D. Mießner
L. Gisotti
M. Kandler
A. Nava
A. Onillon
D. Manfrin
C. Fiorini
A. Gavin
P. Voigt

## Slide 33
Back-Up

## Slide 35: Noise Curve for Detector Module
Electronic noise MPP & MoS
Very compatible in both setups for short shaping times
Slightly higher noise for MoS at higher shaping times
  Looks like increased leakage current (but detector was colder)
  Probably light coming from not perfectly light tight view port
Detector electronic noise almost not effected by nosier MAC-E environment of MOS