## Slide 1: Testing PXD6 - testing plans
Jelena Ninkovic for the HLL team

## Slide 2
Measurements on the first wafers with full DEPFET + THINNING
technology !!!!

## Slide 3: PXD6 Batch 1 wafers
  L1 - lithography dummy but has some of the implantations
S1, W02 – hot wafers with ZMI=600nm
W21 – hot wafer with ZMI=1000nm and Plasma etching of Polysilicon
W11 – Reference wafer (not SOI material)
  Some parts of the wafer have 100nm thinner Oxide then the rest of the wafer

## Slide 4: IR microphotographs
Al lines on n-side through the thick handle and top wafer
Small matrices
Al lines on p-side
Al lines on n-side through the thin top wafer
Edge of the opening window on the p-side

## Slide 5: IR microphotographs
Half ladders

## Slide 6: PXD6 Batch 1
8 (6+2) SOI wafers + 2 reference wafers
3SOI + 1 Ref in the first iteration +1 LithoDummy
128x16 matrices for wire bonding 128x8 matrices for wire bonding 85 standard 8 of them on thin oxide 6 ILC (128x128) 2 of them on thin oxide – (2 best DUTs from TB2009)
Mini matrices (9 of which 2 on thin oxide)
single DEPFETs (24 pairs with varied channel and clear gate geometry)
Some special test structures

## Slide 7: Testing system
Wafer level
  Semiautomatic wafer probing
SUSS PA150 Semiautomatic Probe System
embedded into Keithley 4200
Semiconductor parameterization system

## Slide 8: Depletion voltage
Spread on different wafers
10-15V

## Slide 9: Back side diodes on SOI material
Handle wafer
n+
p+
Al
200nm Oxide

## Slide 10: Back side diodes

## Slide 11: Back-side biasing
Handle wafer
n+
p+
Al
p+
Al
Simplified backside processing
No wire-bonds on thin Si

## Slide 12: Back side biasing

## Slide 13: Breakdowns of dielectrics
Wafer | Poly1 – Poly2 | Al 1- Al2 | P2 – Al1
L1 | -25V +35V | -150V >200V | -20V +120V
W02 | -30V +40V | -125V +135V | -25V +130V
W21 | -25V +35V | -150V +125V | -25V +50V
 Safe operation within the nominal voltage ranges

## Slide 14: Transistors
DEPFET like transistors nicely match the threshold voltage around 0V
High homogeneity over the wafer surface

## Slide 15: Transistors

## Slide 16: Transistors
Flat position

## Slide 17: Transistors

## Slide 18: Transistors – Internal gate potential

## Slide 19: Testing plans
Static measurements - Wafer level measurements:
  Special test devices – qualification of the production quality
    Leakage currents
    Depletion voltage
    Threshold voltage of DEPFET
    Threshold voltage of Clear
    Break down voltages of oxides : Poly1 - Poly2 & Al1 - Al2 & Poly2 – Al2
    Contact resistance
    Sheet resistance
    CV on Caps
    Bulk to Bulk contact
  Matrices pretesting – selection of matrices for further testing
    Leakage current of the back diode
    Clear vs. Source
    Gate vs. Source
    Clear Gate vs. Gate
    Some representative sample of IV curves