## Slide 1: Soft X-ray Detection with JUNGFRAU and LGAD Sensors
2nd Workshop on Silicon Sensors for Radiation Detection and Quantum Applications Lake Starnberg, 13 May 2025
Viktoria Hinger1,*, Anna Bergamaschi1, Ashish Bisht2, Maurizio Boscardin2, Maria Carulla1, Matteo Centis Vignali2, Roberto Dinapoli1, Nuno Duarte3, Khalil Ferjaoui1, Francesco Ficorella2, Erik Fröjdh1, Viveka Gautam1, Dominic Greiffenberg1, Omar Hammad Ali2, Julian Heymes1, Loïc Le Guyader3, Shuqi Li1, Vadym Kedych1, Carlos Lopez-Cuenca1, Davide Mezza1, Konstantinos Moustakas1, Aldo Mozzanica1, Jonathan Mulvey1, Eugenio Paris1, Giovanni Paternoster2, Kirsty A. Paton1, Marco Ramilli3, Elia Razzoli1, Sabina Ronchin2, Andreas Scherz3, Justine Schlappa3, Bernd Schmitt1, Patrick Sieberer1, Saverio Silletta1, Elisabeth Skoropata1, Monica Turcato 3, Hiroki Ueda1, Xiangyu Xie1, Jiaguo Zhang1
1 Paul Scherrer Institut, 2 Fondazione Bruno Kessler, 3 European XFEL GmBH
* viktoria.hinger@psi.ch

## Slide 2: Paul Scherrer Institute
PSI East
Energy Research
Nanotechnology
Neutron Source
SINQ
Swiss Myon Source SµS
Radiopharmacy
Biology
Particle Physics
Proton Therapy
Proton Accelerator
SwissFEL
Swiss Light Source (SLS)

## Slide 3: Photon Science at PSI
Swiss Free Electron Laser (SwissFEL)
Swiss Light Source (SLS)

## Slide 4: Photon Science with Soft X-rays
L-edges of many important metals
  Mn, Fe, Cu …
  Quantum materials
  Spintronics
  Superconductors …
“Water window” and light elements
  C and O K-edges (282 – 543 eV), Na, Mg, …
  Biological samples
  Catalysis
  Chemistry …
Resonant imaging and spectroscopy
e.g. Resonant Inelastic X-ray Scattering (RIXS)

## Slide 5: Problems of Soft X-ray Detectors: Area and Speed
Charge Coupled Device (CCD)
CMOS image sensor
Commercial soft X-ray detectors
Readout speed ~ 1 Hz
Area ~ 2 × 2 cm2
Readout speed ~ 70 Hz
Area ~ 4 × 4 cm2

## Slide 6: Hybrid Detectors are a Viable Solution
Advantages
High frame rate (≥ 2 kHz)
Large dynamic range (120 MeV charge integrating, “unlimited” for photon counters)
Large, customizable area (4 × 8 cm2 single module, tileable)
We can use different sensors to accommodate different X-ray energies
    e.g., HighZ materials for > 20 keV photons
  We can optimize sensor and chip separately!
Sensor
Indium bumps
Readout electronics
To data acquisition system
Chip controls
High voltage to sensor
X-ray
-
+
-
-
-
+
+
+

## Slide 7: Detector Portfolio at PSI
 | Microstrip 1D | Pixel 2D | 
 | 25/50 µm pitch | 75 µm pitch | 25 µm pitch
Single photon counting |  |  | 
Charge integrating |  |  | 
MYTHEN
MATTERHORN
EIGER
GOTTHARD
JUNGFRAU
MÖNCH

## Slide 8: JUNGFRAU – Low Noise, High Dynamic Range
Architecture | Charge-integrating
Technology | 110 nm
Pixel size | 75 × 75 µm2
Counter depth | 14-bit off-chip ADC
Dynamic range | 120 MeV with 3 dynamic gains
Noise (ENC) | 34 e- r.m.s. in high gain, 5 µs integration time
Module size | 4 × 8 cm2, 500k pixel
Readout | 2 × 10 Gbit/s
Frame rate | 2.4 kHz continuous, 250 kHz burst mode (16 memory cells)
Photon rate | 20M 12 keV photons/pixel/s
Applications | XFELs and high-flux synchrotron experiments
Architecture | Charge-integrating
Technology | 110 nm
Pixel size | 75 × 75 µm2
Counter depth | 14-bit off-chip ADC
Dynamic range | 120 MeV with 3 dynamic gains
Noise (ENC) | 34 e- r.m.s. in high gain, 5 µs integration time
Module size | 
Readout | 2 × 10 Gbit/s
Frame rate | 2.4 kHz continuous, 250 kHz burst mode (16 memory cells)
Photon rate | 20M 12 keV photons/pixel/s
Applications | XFELs and high-flux synchrotron experiments

## Slide 9: Limitations of Hybrid Detectors in the Soft X-ray Range
Low quantum efficiency (QE)
Recombination in the n+ layer and at the Si/Al interface
Absorption in the surface layers
Low signal-to-noise ratio (SNR)
SNR > 5 required for single-photon resolution
With standard Si sensors SNR 1.6 at 200 eV
Single-photon peak at 800 eV
with typical JUNGFRAU noise
V. Hinger et al. 2022 JINST 17 C09027
Thin Entrance Window (TEW)
Reduce losses at the entrance window
Low Gain Avalanche Diode (LGAD)
Internal signal amplification

## Slide 10: Thin Entrance Window (TEW) Technology
Standard sensor
Sensor with shallow doping
Optimized TEW (standard passivation)
Optimized TEW (thin passivation)
M. Carulla et al. 2024 Sensors 24, 942
To improve quantum efficiency (QE)

## Slide 11: LGAD Sensors for Soft X-ray Detection
To improve signal-to-noise ratio (SNR)
Soft X-ray detection capability
Measured at SLS SIM beamline with iLGAD bonded to MÖNCH (25 µm pixel)
Inverse design (iLGAD)
To achieve 100% fill factor
900 eV
390 eV

## Slide 12: iLGAD Double-Peak Spectral Response
900 eV
Detector output [arb.unit]
Me
900 eV
500 eV
250 eV
Mh
n+
gain layer
X-rays
depth
gain
attenuation
Courtesy of Jiaguo Zhang
Soft X-ray detection capability
Measured at SLS SIM beamline with iLGAD bonded to MÖNCH (25 µm pixel)
900 eV
390 eV
absorption
beyond gain layer Me
absorption
before gain layer Mh

## Slide 13: iLGAD Double-Peak Spectral Response
900 eV
Detector output [arb.unit]
Me
900 eV
500 eV
250 eV
Mh
n+
gain layer
X-rays
depth
gain
attenuation
Courtesy of Jiaguo Zhang
SNR > 5 for single photons down to 400 eV
To reach 200 eV the gain needs to be increased by at least a factor 2!
900 eV
390 eV
absorption
beyond gain layer Me
absorption
before gain layer Mh
absorption beyond gain layer
SNR iLGAD
absorption before gain layer
SNR standard Si sensor

## Slide 14: Combining JUNGFRAU with iLGAD Sensors
First prototype | 
Sensor pixel size | 75 × 75 µm2
Module size | 2 × 2 cm2, 56k pixel
ASIC version | JUNGFRAU 1.1
Noise (ENC) | 5.5 e- r.m.s., 5 µs integration @ -22°C
Bad pixels | 0.8 – 1.5%
Leakage current | 5 pA/pixel @ -22°C
Multiplication factor | 11 @ -22°C
First prototype | 
Sensor pixel size | 75 × 75 µm2
Module size | 
ASIC version | JUNGFRAU 1.1
Noise (ENC) | 5.5 e- r.m.s., 5 µs integration @ -22°C
Bad pixels | 0.8 – 1.5%
Leakage current | 5 pA/pixel @ -22°C
Multiplication factor | 11 @ -22°C
V. Hinger et al. 2024 Front. Phys. 12

## Slide 15
Our Target Application

## Slide 16: Resonant Inelastic X-ray Scattering (RIXS)
Scientific goal
Access very specific, complex electronic-structural information and dynamics
Charge, spin, orbital, and lattice excitations
Quantum materials: spintronics, superconductors
Method
Scan over electronic resonance
Measure transferred energy and momentum

## Slide 17: Resonant Inelastic X-ray Scattering (RIXS)
Scientific goal
Access very specific, complex electronic-structural information and dynamics
Charge, spin, orbital, and lattice excitations
Quantum materials: spintronics, superconductors
Method
Scan over electronic resonance
Measure transferred energy and momentum
Characteristics
Energy resolution ↔ spatial resolution
“Photon-hungry”
Single shot detection at high repetition rates
Multi-dimensional scans
Statistics
Energy resolution
Elastic line (930 eV)
Courtesy of Nuno Duarte
Laser pump

## Slide 18: First JUNGFRAU for Soft X-ray RIXS
Sensor pixel size | 225 × 25 µm2
Module size | 4 × 4 cm2, 240k pixel
ASIC version | JUNGFRAU 1.0 (16 memory cells)
Noise (ENC) | 9 e- r.m.s., 5 µs integration @ -20°C
Sensor pixel size | 225 × 25 µm2
Module size | 
ASIC version | JUNGFRAU 1.0 (16 memory cells)
Noise (ENC) | 9 e- r.m.s., 5 µs integration @ -20°C
effective ENC (eV)
Native pixel resolution
Interpolated
75 µm
75 µm
Implant
Metal
X-ray photon
25 µm
Interpolation direction
  Pixel pitch must be comparable to the spread of the charge cloud!
  Sensor with rectangular pixels
225 µm

## Slide 19: Using Charge Sharing to Perform Position Interpolation
To achieve high spatial resolution µm
25 µm pixels
2.5 µm bins
MÖNCH-iLGAD module (25 × 25 µm2 pixels)
1 cm
1 cm
Courtesy of Jiaguo Zhang
Measured at SLS POLLUX beamline
Photon energy: 500 eV
Estimated position resolution: 2.7 µm

## Slide 20: Test at European XFEL Heisenberg RIXS Spectrometer
Marana-X sCMOS
JUNGFRAU iLGAD
EuXFEL pulse structure
Pixel size | 25 × 225 µm2
Frame rate | 250 kHz burst
Pixel size | 11 × 11 µm2
Frame rate | 70 Hz
Pixel size | 11 × 11 µm2
Frame rate | 
Exit arm ( 5 m)

## Slide 21: Test at European XFEL Heisenberg RIXS Spectrometer
Marana-X sCMOS
Pixel size | 25 × 225 µm2
Frame rate | 250 kHz burst
Pixel size | 11 × 11 µm2
Frame rate | 70 Hz
Pixel size | 11 × 11 µm2
Frame rate | 
Exit arm ( 5 m)
Elastic scattering on carbon tape
Courtesy of Nuno Duarte
Marana-X
JUNGFRAU iLGAD

## Slide 22: Test at European XFEL Heisenberg RIXS Spectrometer
Marana-X sCMOS
JUNGFRAU iLGAD
Pixel size | 25 × 225 µm2
Frame rate | 250 kHz burst
Pixel size | 11 × 11 µm2
Frame rate | 70 Hz
Pixel size | 11 × 11 µm2
Frame rate | 
First ever RIXS spectra alternating pumped (P), unpumped (UP) in a single pulse train (@ 930 eV)
Exit arm ( 5 m)

## Slide 23: Test at European XFEL Heisenberg RIXS Spectrometer
Marana-X sCMOS
JUNGFRAU iLGAD
Pixel size | 25 × 225 µm2
Frame rate | 250 kHz burst
Pixel size | 11 × 11 µm2
Frame rate | 70 Hz
Pixel size | 11 × 11 µm2
Frame rate | 
Decrease of photons emitted from sample within a train, affected by FEL intensity
Exit arm ( 5 m)

## Slide 24: Summary & Outlook
Achievements
QE > 80% @ 250 eV with Thin Entrance Window
SNR > 5 for single photons down to 400 eV with iLGADs
First promising results at RIXS experiments
Current challenges
Efficiency at low photon energies and shallow incidence angles
Absorption depth dependent multiplication factor due to inverse LGAD design
Goals for the new iLGAD batch (available 2026)
Reduced oxide thickness
Increased gain
Shallower gain layer
Detecting soft X-rays with JUNGFRAU and iLGAD sensors
4 cm
4 cm

## Slide 25: Brief Advertisement: X-Ray Sensor Workshop Upcoming!
See you there!

## Slide 26: Acknowledgements
European XFEL
Nuno Duarte, Loïc Le Guyader, Marco Ramilli, Andreas Scherz, Justine Schlappa, Monica Turcato
PSI Center for Photon Science Detector Group
Front: Roberto Dinapoli, Anna Bergamaschi, Shuqi Li, Dhanya Thattil, Jiaguo Zhang, Alice Mazzoleni, Viveka Gautam, Davide Mezza, Thomas King, Bernd Schmitt Back: Konstantinos Moustakas, Christian Ruder, Vadym Kedych, Jonathan Mulvey, Martin Müller, Patrick Sieberer, Saverio Silletta, Khalil Ferajoui, Viktoria Hinger, Martin Brückner, Aldo Mozzanica, Erik Fröjdh, Maria Carulla, Julian Heymes, Xiangyu Xie, Carlos Lopez-Cuenca Not on photo: Simon Ebner, Shqipe Hasanaj, Dominic Greiffenberg, Kirsty A. Paton
SwissFEL Furka
Eugenio Paris, Elia Razzoli, Elisabeth Skoropata, Hiroki Ueda
Fondazione Bruno Kessler
Ashish Bisht, Maurizio Boscardin, Matteo Centis Vignali, Francesco Ficorella, Omar Hammad Ali, Giovanni Paternoster, Sabina Ronchin

## Slide 27
Backup

## Slide 28: Thin Entrance Window Radiation Hardness
Diodes measured at SOLEIL synchrotron
Courtesy of Jiaguo Zhang

## Slide 29: JUNGFRAU-iLGAD Effective Noise

## Slide 30: Interpolating with Rectangular Pixels
Clustering to reconstruct photon hit
 Attribute shared charge to central pixel
Eta function
 Accounts for non-linear behaviour of charge sharing
 From flat-field: construct histogram from charge fraction of each pixel
 Integral gives the interpolation curve (lookup table)
 | 0.97 | 
 | 0.53 | 
X-ray photon
12 µm

## Slide 31: Validating the Position Resolution with 1-µm Size Beam
2.5 µm steps
Estimated resolution
 2.1 µm
 2.3 µm
 2.6 µm
 2.7 µm
 2.2 µm
 2.3 µm
 2.3 µm
Interpolation direction
75 µm
75 µm
25 µm
225 µm
distribution for different scan positions
Implant
Metal
1 keV
@ SLS Pollux beamline

## Slide 32: Test at SwissFEL Furka
Exit arm ( 5 m)
Detector chamber
Grating chamber
Experimental chamber
X-rays scattered from sample
Image rate aligned with SwissFEL (100 Hz) allows efficient jitter correction for time-resolved measurements
Good performance at Cu L-edge (~ 930 eV), but low efficiency at O K-edge (~ 530 eV) because of shallow incidence angle
Pixel Y
Pixel X
# photons
Pixel Y
FWHM 62 µm
Elastic scattering off carbon tape @ 932 eV
No interpolation,
no curvature correction
Preliminary
Preliminary

## Slide 33: Test at SwissFEL Furka
Exit arm ( 5 m)
Detector chamber
Grating chamber
Experimental chamber
X-rays scattered from sample
Image rate aligned with SwissFEL (100 Hz) allows efficient jitter correction for time-resolved measurements
Good performance at Cu L-edge (~ 930 eV), but low efficiency at O K-edge (~ 530 eV) because of shallow incidence angle
5 effective noise
Hole peak
Electron peak (500 eV)
Not detected

## Slide 34: Optimizations Planned with Upcoming iLGAD R&D Batch
Electron peak (500 eV)
5 (new batch)
5 effective noise (today)
Hole peak
Hole peak
Reduce oxide thickness
Improve quantum efficiency
Reduce thickness of multiplication layer
Improve Mh/Me
Increase iLGAD gain
Improve signal-to-noise ratio
@ 500 eV
Combined detection efficiency (500 eV photons, 30° angle) 10% (today)
Combined detection efficiency (500 eV photons, 30° angle) 10% (today) > 80% (new iLGAD batch)