## Slide 1: Nonlinear Amplification revisited Insights and modelling
2nd Workshop on Silicon Sensors for Radiation Detection
and Quantum Applications
Berg, 12.05.2025

## Slide 2: EDET DH80K
1 MPixel direct hit electron detector with 80 kHz framerate
J. Treis, L. Andricek, M. Hensel, C. Koffmane, J. Ninkovic, M. Polovykh, M. Predikaka, E. Prinker, R. Richter, M. Rizwan, T. Selle, A. Vostrukhin
MPG Semiconductor Laboratory
S. Epp, D. Gitaric
MPSD Hamburg
I. Peric
KIT Karlsruhe
T. Hemperek, H. Krüger
Bonn University
Applying the principle of stroboscopic imaging in the TEM world to enable measurements with extremely high time resolution

## Slide 3: Contents
Nonlinear amplification – why?
Accuracy functions
The DEPFET and NLA
The many gq-s of the DEPFET
Calculating DEPFET gq
Summary & Outlook

## Slide 4: Requirements
Stroboscopic imaging in TEM world is challenging:
Fast imaging → Direct electron detection
Real space imaging → Large sensitive area
High intensity → High dynamic range
High contrast range → Signal compression in sensor
Direct electron detection → Thin substrate
High pulse frequency → High framerate
"Grey scale" image → No data reduction possible
Small volume → Extremely high integration density

## Slide 5: Requirements
Dynamic range
  Single primary electron sensitivity
  Primary electron @ 300 keV on 50 µm Si
  → Distinguish 8k e- from noise
  50 (better 100) primaries per pixel to provide enough contrast
  → 800k signal electrons to be stored in pixel
Operation in vacuum:
  Small volume / extremely compact setup
  Cooling / thermal stabilization difficult
Multiple scattering:
  Thin sensitive detector substrate
  No support layer
  Highly efficient beam dump for traversing electons
  Advanced thermal support
Data rate
  Max. burst frequency 100 Hz ( 50 frames each)
  Digitization with 8 bit resolution
    Tile module data rate 1.28 GByte/s gross
    Total data rate ~ 6.4 GByte/s gross
    Data reduction / zero suppression difficult
Translation:
System noise must be significantly lower than minimum signal amplitude
Translation:
Amplitude for maximum stimulus must be
smaller than or equal to maximum amplitude

## Slide 6: Accuracy function
Needed:
Model tool to investigate interdependence between:
  Maximum input signal
  Maximum output amplitude
  Nature of signal
  System gain
  System noise
  Noise threshold
  Minimum detectable signal
... using some simplifications.
Accuracy function
Detector is black box converting a stimulus into an amplitude
Stimulus has variations depending on its nature
Amplitude is limited () and adds a (constant) noise contribution
Stimulus can be recovered using an effective calibration function
Recovered stimulus has variation
Accuracy function describes the error on the reconstructed stimulus as a function of the original stimulus:

## Slide 7: Accuracy function
Stimulus:
Total stimulus is sum of N individual quanta (electrons / photons):
Monoenergetic, varying around a common mean stimulus:
For counting applications:
is the signal to be retrieved.
Error given by Fano and „Landau“

## Slide 8: Accuracy function
Finding:
Any unspecified gain function
Yields a generalized accuracy function
Needed:
First derivative of the inverse wrt. A
Partial derivatives wrt. S and N
„Valid“ gain function must be:
Invertible, continuous
Derivative of the inverse must exist
Must be (piecewise) differentiable with respect to S and N
Offset, AKA
Pedestal, AKA
Black level ....

## Slide 9: Accuracy function
Example 1: Linear gain
Linear gain function „ansatz“:
Linear gain coefficient (from boundary conditions):
Accuracy function:
Using the amplifier dynamic range to noise ratio (ADNR):

## Slide 10: Accuracy function
Example 2: Semilinear gain
Semilinear gain function ansatz:
Semilinear gain coefficient (from boundary conditions):
Accuracy function:

## Slide 11: Accuracy function
Other quantities to be calculated:
Low energy limit (minimum stimulus distinguishable from noise) with
Dynamic range:
Single quantum sensitivity limit
: noise discrimination factor (application specific)

## Slide 12: Accuracy function
 | Linear gain function | Semilinear gain function
Gain function |  | 
Gain coefficient |  | 
Effective calibration function |  | 
Accuracy function |  | 
Low energy limit |  | 
Dynamic range |  | 
Single quantum sensitivity limit |  | 
Dynamic range to noise ratio:
Note:
For most applications, it can be taken for granted that

## Slide 13: Accuracy function
Comparison:
Stimulus error
Counting error (Poisson statistics)
Error due to amplifier noise
Linear
gain function
Semilinear
gain function
Break-even at
For smaller , the accuracy for the semilinear gain function is higher
For larger , the accuracy for the linear gain function is higher
For , the ASF for the linear GF stays constant, while the ASF for the semilinear GF drops to 0.
Break-even at 25
This term also shows up in every gain function
scaling factor (ASF)

## Slide 14: Accuracy function
Conclusion:
Choosing a gain function is about distributing output dynamic range to accuracy of retrieved stimulus
Zero-sum situation: Achieving high accuracy in one stimulus region yields lower accuracy in other regions
Accuracy function has Poisson error as built-in limitation
There is no „generic“, or optimal, gain function. It‘s all about user requirements and priorities.
Parameters:
  Input stimulus and input stimulus „error“
  Maximum requested stimulus
  Minimum stimulus to be detected
  Maximum output Amplitude
  ADNR
  „Tertilinear“:
  „Sine “:

## Slide 15: DEPFET and nonlinear amPlification

## Slide 16: DEPFET and NLA
DEPFET integrated amplifier
p-channel FET on depleted n-bulk
Signal charge collected in potential minimum below FET channel
"Internal gate", low capacitance & noise
FET current modulation ≥ 300 pA/el.
Reset via clearFET
Charge storage
Readout on demand
Rolling shutter mode
EDET pixels:
Dynamic range problem
Implement signal compression in pixel
Overflow to source to tailor dynamic range

## Slide 17: DEPFET and NLA
Signal charge (Me-)
Drain current (100 µA)
Response curve (VDS = -5 V)
Primary OF "Kink"
CHC limit
Secondary
"Kink„?
Signal charge (Me-)
Signal charge (Me-)
Charge distribution (VGS = -3 V, VDS = -5 V)
Response curve, or Charge-current (C2)-curve
OV1
OV2

## Slide 18: The many of the DEPFET
Simple voltage depending capacitive model w/ 3 zones
Constant „generic“ default of IG
Only charge in „internal gate“ contributes to
Here, 3 steps in can be expected!

## Slide 19: The many of the DEPFET
Common definition of
But behold the quantity
This quantity is also often used because of
Same dimension, but the quantities are identical only in small signal mode or for devices w/o NLA.
„Differential“ : (Local) slope of the C2-curve!
„Integral“ : average slope for charge interval between and
May cause confusion if mistaken!

## Slide 20: The many of the DEPFET
Differential and integral can be converted into each other, as
and
Thus, the net C2-curves can be calculated by
or

## Slide 21: The many of the DEPFET
Because of
The ratio of the is linked to the area ratios:
Already something, but model is too simple
Shortcomings of the model:
  Gives ratios for s, not values
  Overflow charge/voltage points are not given
  Inductance effects are ignored
Need to use more sophisticated and fundamental approach

## Slide 22: Calculting DEPFET
Approach:
Use actual space charge density in storage region (as simulated) to calculate both and
Modify and extrude (simulated) charge cloud to investigate different layouts
Later: Use (technology dependent) parametrization of charge in internal gate to model arbitrary layouts

## Slide 23: Calculting DEPFET
Very generic model „transfers“ internal gate charge to the external gate to calculate :
Conversion factor is the capacitance of external gate:
„Loss“ factor quantifies the fraction of the internal gate charge which is effective, i.e. „seen“, by the DEPFET channel
For the operation of the DEPFET in small signal mode (i.e. constant ), a constant factor is assumed
For large charge quantities or devices with overflow regions, the factor has to be made a function of the charge:
The factor evolves from simple „loss factor“ to charge depending „influence effectivity factor“ (IEF)
How to calculate the IEF?

## Slide 24: Calculting DEPFET
Calculate charge density at location of DEPFET channel as induced by charge in every voxel in the respective regions of the storage
Quotient of simulated charge vs. induced charge is inductance efficiency
Apply Shockley-Ramo Theorem to charge cloud in storage region
Step 1: Calculate the induced charge in the channel
Step 2: Calculate the IEF
Weighting potential of channel region

## Slide 25: Calculting DEPFET
The weighting potential usually calculated numerically
Here, simplified (e.g. neglect influence of backside, surface „flat“ etc.) analytical solution is used
Corresponds to simple „mirror charge over conductive plate“ problem
Weighting potential
This yields a closed algebraic epression
Error estimation yields error in the 1% range for the thinnest detector values

## Slide 26: Calculting DEPFET
Very nice match with few parameters!
Compare model calulation with numerical simulations
Nice agreement
Few average parameters (mobility etc...)
Geometry parameters
Gate-Source voltage...
To a large degree, the gq dependence seems to be an „electrostatic“ phenomenon

## Slide 27: Calculting DEPFET
1:
0B:
2:
3:
Different designs based on extruded simulation results
Distance of overflow 2 to internal gate becomes smaller
Overflow 1 brim introduced
Area ratios between overflow 1 and overflow 2 regions are very similar
External gate position and charge distribution in internal gate are not modified
Interesting parameters:
  Device capacitance values for the various regions
  Overflow thresholds
  gq as a function of charge

## Slide 28: Calculting DEPFET
EDET pixel design:
Situation similar to case 3 in terms of Gate - OV1 gap
Much wider brim (9 -10 µm)
Area ratio changed wrt. reference design
Probably mismatch in gate width
Model:
Design:

## Slide 29: Calculting DEPFET
Next step: Calculate Charge-current (C2)-curve

## Slide 30: Calculting DEPFET
The can be used to calculate the C2-curves:
Charge-current (C2)-curve is experimentally relevant DEPFET characteristics
Gives overall current in the pixel as a function of the total signal charge in the pixel
If the detector is properly calibrated, the C2 curve can be used to determine the verall signal charge.
„Total“ C2-curve and net C2-curve (w/o )

## Slide 31: Calculting DEPFET
Net C2-curve of OSCAR simulation
Net C2-curve calculated for EDET-like design

## Slide 32: Calculting DEPFET
Because of
the is linked to the capacitance!
Example: Design 0B
Steps reflect the activation of the various overflow regions
Kinks are (as expected) much more pronounced as in the integral plot

## Slide 33: Construction of Gain function for DEPFET
1st step: Define parameters
Input parameters: , , type of
Choose gain function (may be linear interpolation)
Define „gain“ coefficient(s)
Determine required
Fix kink points according to preference / requirement
2nd step: Convert gain function to layout
Input: Charge density function (simulated or parameterized) alongside activation threshold voltages
Calculate internal gate size according to required primary .
Set dimensions (if possible) such that threshold voltage meets requested first kink point
Set area of overflow 1 such, that secondary requirement is met.
Check if kink point requirement is met (variation here is difficult, if no modification fo technology possible)
Set area of overflow 2 such that tertiary gq requirement is met
Check, if requirement is met
Caveats:
Many interdependencies,
Limited design options

## Slide 34: Summary
Accuracy functions are a useful tool for the prediction of performance of a detector system in terms of measurement accuracy
Can be used to define an optimum gain function for an application according to the respective requirements
Mainly useful for counting applications, significance for spectroscopy is limited
The Shockley-Ramo theorem can be used to predict the to be expected for a given layout
Modelling method was developed and tested
Increases understanding of amplification mechanism in DEPFET
Helps to narrow down layout options prior to simulation for design optimized for a specific application
Insights gained are significant also for other DEPFET applications

## Slide 35: Outlook
Create parametrization for charge density function
Optimize design
Verify optimized designs by actual 3D-Simulation
Investigate and quantify effect of inductance on operational parameters for other DEPFET applications (RNDR, Inifinipix, Quadropix)
Use insights gained to study the amplification mechanism in the DEPFET channel in detail.
Learn about the effect of charge distribution in the channel on charge amplification
Optimized positioning of internal gate
Imporved performance

## Slide 36: Questions?