## Slide 1: Fast Small Pixel - The next Generation (FSP TNG)

## Slide 2: MAXYMUS/MYSTIIC @ Bessy II
BESSY II, Helmholtz-Zentrum, Berlin
*MAgnetic X-raY Microscope
with UHV Spectroscopy

## Slide 3: Ptychography
record n pictures
move sample in x and y
do this for m positions
reconstruct sample from m x n pictures
5 nm half-pitch resolution aperture limited
200 nm
Ptychographic reconstrution of
commercial Fresnel Zone Plate

## Slide 4: ptychography
Numerical aperture (NA)
Reconstractions with different NA

## Slide 5: Concept

## Slide 6: pnCCD
  definition of potential pockets by differently reverse-biased diodes
  charge transport by periodic clocking of shift registers
  column-parallel readout  high frame rate (5 msec @ 200 pixel)
  integrated 1st FET (1 / column)  low noise (2el. ENC)
  backside illuminated, fully depleted  high quantum efficiency
  Applications
    X-ray imaging & spectroscopy
    optical light imaging
  format ~ cm² … wafer scale
  thickness 450 µm
  pixel size 36 … 150 µm

## Slide 7: FSP TNG
build a camera around a
  1 MPix pnCCD
  36 x 36 µm² Pixels
  framestore area
  split frame readout
develop:
thermo-mechanics
electronics
soft-and firmware for operation
(dummy) Sensor + Hybrid PCBs

## Slide 8: FSP TNG - concept
sensor + hybrids inside vacuum
IIC + mezanine cards outside vacuum
  Control over Mercury XU1
  1 GB – slow control from PC
  10 GB Ethernet – datatransfer XU1 to Storage

## Slide 9: Mechanics

## Slide 10: Mechanics – camera Head
water cooling
cable fixtures
xyz stage
camera head
MAXYMUS mechanics

## Slide 11: Vacuum FeedthrougH
Cable bracket (x4)
vacuum flange (iso-f)
feedthrough pcb
vacuum flange (custom)
IIC (Instrument Interface Card)
Al-Housing
ISO-K to CF adapter

## Slide 12: Vacuum FeedthrougH- cables
Concept: PCB clamped between 2 costum flanges
4 identical cables connecting to Camera Head

## Slide 13: The „Brain“ Firmware

## Slide 14: FSP TNG Firmware
Mercury XU1 SoC Board
XCILINX Zynq SoC

## Slide 15: FSP TNG Firmware
divide tasks between
      processing system
        on-chip linux
        drivers for PL
        houskeeping software
      processing logic
        communication interfaces
        sequencer
        data reduction

## Slide 16: FSP TNG Firmware
accumulation of frames
for every (ptychography) position
stack multiple frames inside FPGA
send only accumulated data to storage
increase dynamic range
reduce data rate

## Slide 17: Current state

## Slide 18: Hardware test
CCD test on ATG
Automated pre-test of CCDs
phi-registers, MOS, bulk tested against each other
9 devices tested (including dummies)
3 devices selected for further assembly

## Slide 19: Mechanics check at Maxymus

## Slide 20: Electronics testing
Electronic components: FPM tested IIC 2 PCBs tested and calibrated 2 more populated HIC 4 PCBs tested and calibrated ADC card 2 cards under test Phi driver several cards tested Hybrids populated with ASICs - debugging and testing

## Slide 21: Next steps
finish test of electronic components
assembly of camera head
glue pnCCD
mechanical assembly
bonding
test of system inside of vacuum chamber
first time combination of all components

## Slide 22: Thanks for listening