## Slide 1: DCDpp Transfer curves
Manfred Valentan, Christian Koffmane, Eduard Prinker

## Slide 2: DCDpp Test System on “hybrid4”
Switcher
for Clear and Gate Control
2nd Switcher
Not used for Belle-II type PXD6 Design
PXD6 Belle-II DEPFET
DCDpp
Read-out Chip
DCD-RO
Line Driver and Buffer to FPGA

## Slide 3: Overlay, pedestals
Measurement procedure:
Calibrate internal source with SMU
Select DCD channel [0,255]
Step through current values [0,384]
Record 100 samples, save noise
Analysis: use single samples, not the mean
-
-
σ
-

## Slide 4: Observables – Gain and noise
k = gain

## Slide 5: Gain, noise
Channel 121
σ
σ

## Slide 6: Channel 121

## Slide 7: Observables – Dynamic range

## Slide 8: Dynamic range
Channel 128
σ
σ
σ

## Slide 9: Channel 128

## Slide 10: Observables – Missing codes

## Slide 11: Missing codes
Note: only consider 5 or more consecutive missing codes

## Slide 12: Missing codes
Channel 57

## Slide 13: Channel 57

## Slide 14: Influence on position measurement
Study by Peter Kodyš
3 bit missing code (8 ADUs): resolution worsens by 0.9 µm (10%)
Prerequisites:
  The affected channel is part of a cluster -> what are the odds?
  Primary signal below 20 ADU -> maybe even below threshold...
Fewer missing codes and higher signal -> smaller effect
Conclusion: missing codes are ugly, but don‘t harm the detector performance

## Slide 15: Summary
Gain:
  Channel average 73 nA/ADU
  Periodicity with half column
  Curved shape
Noise:
  Channel average 0.4 ADU
  Maximum 2.2 ADU (mean of a single channel)
  4 channels with oscillations
Dynamic Range:
  Channel average 7.8 bit
  All channels show nonlinearity at low end
  Inner four columns show nonlinearity also at high end
Missing codes
  20 channels (7.8%)
  Up to 14 adjacent missing codes
  Most zones start at 64
  No dependence on geometry

## Slide 16: What‘s next
Same measurement on EMCM
  Speed up measurement
  Try for different DCDs
  Try for different supply voltages and settings
  Study effect of neighbouring chips
Refine algorithms
  Make them fit for quality assurance
  Store results in database for later reference
DCD feature request: Finer internal current source
  However, only „nice to have“ for detailed lab tests

## Slide 17
Thank you
for
your attention!

## Slide 18
Backup

## Slide 19: How to find the linear range
Smoothen the transfer curve by “moving average”
Linear fit between fixed lower limit and variable upper limit -> plot Chi2/ndf
Linear fit between fixed upper limit and variable lower limit -> second half
First derivative of Chi2/ndf curve
Apply cut, disregarding small spikes

## Slide 20: Static ADC Parameters (expressed in LSB-units or %FSR)
quantization error: ½ LSB resolution e.g. 8 bit-resolution = 28 possible digital codes 
offset error: D of of actual from perfect transfer function at point of zero
gain error: D of last step‘s midpoint of actual from perfect transfer function
gain: slope of a least mean squared straight line [nA/ADU]
noise: measured in ADUs
(not shown) INL(integral non-linearity) = deviation of digital output code from its ideal location
(not shown) DNL(differential non-linearity) = deviation of digital output code width from ideal width (=1LSB)

## Slide 21: Ideal transfer function of 3-bit ADC
1LSB = 1 ADU = full analog range
Step width = smallest resolution
Uniform staircase with ideally infinite steps:
  tread represents a particular digital output code
  riser represents a transistion between adjacent codes
Transition not sharply defined; resembles more a probability function with slowly increasing input voltage ADC converts more and more frequently to next code
Quantization process introduces error (triangles)

## Slide 22: DNL – missing code
small change in voltages causes a change in result of 2 codes without intermediate code
with VD analog value of digital output code

## Slide 23: DNL – differential non-linearity: Precision

## Slide 24: DNL - monotonicity
lower conversion result for a higher input voltage

## Slide 25: INL – integral non-linearity: Curvature
ideal transfer curve = linear least mean square fit
100 measurements at each analog input step
0 < D < 2N-1

## Slide 26: INL/DNL example

## Slide 27: INL, DNL

## Slide 28
DEPFET sensors at Belle II
Peter Kodyš, 23 October 2014

## Slide 29
DEPFET sensors cases in study (1)
Normal response
Shift one level in one step up
Shift one level in two steps down
Shift two levels in two steps up/down
Shift two levels in one step up
Every second step shift up
Doubled one ADU
from 200 el. to 400 el.
1
2
3
4
5
6
seeds
Peter Kodyš, 23 October 2014

## Slide 30
DEPFET sensors cases in study (2)
Step down from 7 and next 8
7
8
9
10
11
12
seeds
Peter Kodyš, 23 October 2014
Step down from 12 and next 8
Step down from 24 and next 8
Step down from 35 and next 8
13
14
Step down from 7 and next 4
Step down from 12 and next 4
Step down from 24 and next 4
Step down from 35 and next 4

## Slide 31
1
6
4
5
3
2
seeds: red, cluster charge: green, residual plots
DEPFET sensors cases in study – results (1)
Peter Kodyš, 23 October 2014

## Slide 32
seeds: red, cluster charge: green, residual plots
DEPFET sensors cases in study – results (2)
Peter Kodyš, 23 October 2014
7
8
9
10
11
12
13
14

## Slide 33
1
seeds: red, cluster charge: green, residual plots
DEPFET sensors cases in study - extremes
Peter Kodyš, 23 October 2014
Normal response
RMS 7.6 x 8.6 mm
Cut all digits over 25 to 25 ADU
RMS 7.2 x 8.2 mm
Binary readout
RMS 9.3 x 10.1 mm

## Slide 34
Residuals was difference between true position and reconstructed position
Accumulation of ADU up to 8 up (3 bits error) affect resolution in less 0.9 mm (10 %) for signal less 20, higher signal not (case 7, 8).
Accumulation of ADU up to 4 up (2 bits error) affect resolution in less 0.2 mm (3 %) for signal less 20, higher signal not (case 11, 12).
Rest of cases and residual plot RMS varied in less 0.1 mm in u and v directions
There is no real influence to average residual plots and for resolution
Important note: resolution affection is only for clusters containing affected digit. E.g. if there is 5% of affected digits, there is 5 % contribution of adding error on average error RMS.
Extremes:
  Cut down in 25 ADU have no influence to residual plots and for resolution
  Binary readout gives in average about 2 mm worst resolution (20 %)
Presenting analysis was prepare for 50 x 55 mm DEPFET pixels we are not expecting changes for other pixel size prepared for Belle II.
DEPFET sensors cases in study - conclusion
Peter Kodyš, 23 October 2014