## Slide 1: DCD Specifications
DCD Specifications
Size: 3.2mm x 5mm
Number of channels: 256
Radiation tolerance:20 MRad
Power consumption: less than about 1W
Resolution >10 bit – present chip: 2 bit DAC and 8-bit ADC
PSRR: power supply may differ by 100mV from chip to chip
PSRR (AC): to be tested
Noise: about 100e (ENC)
INL: not so critical
What problems remain to meet the specifications
Some number of ADCs show long missing coded – influences INL
These channels are sensitive for RefIn value (need to be defined as +-50mV)
Changes prior to a subsequent (final) submission
Try to resize several transistors in layout to improve matching
Change JTAG TDI sampling TCK edge
Improve the granularity of both internal test signal sources
Implement some kind of fast parallel sampling of ADC outputs – may allow testing of all ADCs with fewer needles (this can be done even now via JTAG – however it is not optimal because the JTAG clock tree is too slow and has a large skew)

## Slide 2: Switcher Specifications
Switcher Specifications
Size: Should fit to the bump pad pattern – size of the chip may vary
Number of channels: 32
Radiation tolerance: 20 MRad
Power consumption: AC power consumption should dominate
Switching speed: clear rise times should be about 10ns – 20ns for 120pF (presently ~20ns)
Voltage ranges: Amplitude up to 20V, High level: up to 30V (for clear), Low level: down to -10V (for gate)
What problems remain to meet the specifications
The CLEAR switching speed is slightly slower than simulated (now 20ns for 150pF load), however to my opinion DEPFET may work even with present raise times. More measurements are needed. Rise time increase is quite easy to implement – larger output transistors
Implement some kind of fast parallel sampling of switcher outputs

## Slide 3: Switcher Tests
Test plan for Switcher
Important: should we implement extra test pads which are connected to JTAG and used only for testing: Motivation 1: testing of bump bond pads may damage them, motivation 2: test pads can be probed by standard needles
Test JTAG (ID readout, readback) (if it does not work, do not use the chip)
Test the operation with a suitable pattern for normal and gated mode readout, measure all outputs – implement some kind of pattern recognition in FPGA (needs HV-LV converters) (only if pattern is correct in all the channels use the chip) (Alternative: implement HV-LV converter on chip – load the data to JTAG chain)
Use nominal voltages in the tests (Using of lower voltages may be ok to be discussed)

## Slide 4: DCD Tests
Test plan for DCD
Important: should we implement extra test pads which are connected to JTAG and used only for testing: motivation 1: testing of bump bond pads may damage them, motivation 2: test pads can be probed by standard needles
Test JTAG (ID readout, read-back, all digital IOs) (if it does not work, do not use the chip)
Measure all ADC characteristics (every input 10 times) – proposed by Christian Kreidl: measure the ADCs in parallel (8 in parallel)
Comment: present needle card system does measurements at ~100MHz – we need to improve the system to work at 320Mhz, one possibility: parallel sample of the digital signals are read in series
Repeat the measurement for 2 RefIn voltages
Estimation of the test time: (clocking out of all 64 outputs: 640ns)
Number of points 2 (RefIn voltages) x 500 input current points x 10 repetitions x 32 ADCs x 640ns ~ 200 ms if FPGA does the test automatically – data amount 2.56 MByte/chip: goal test one chip in 10 seconds
Store the data as data base for offline analysis
Offline analysis: make the score-table of all chips (e.g. by number of good ADCs) and start using from the best ones