## Slide 1
PXD Ladder Assembly
- Conceptual design of jigs -

## Slide 2: The Mission
Join two modules to one ladder
Modules:
  fully assembled half-ladder
  electrically fully functional
  Interconnect to rest of world through kapton
  fully tested
Ladders:
  Ready to be installed in PXD support structure

## Slide 3: 1st trials 2010
:- thinned silicon dummies, 50 µm “sensitive”
:- fully manual, just to test feasibility

## Slide 4: Conclusions then: in principle okay..
:- joint between two modules very stable!
:- alignment and procedure need improvement
:- still, can do it better …
:- but need dedicated jigs for this job!

## Slide 5: Requirements
The jigs should allow:
:- safe storage and transport of modules and ladders
:- support for SMD mounting
:- module testing on the probe stations
:- support for kapton attachment (soldering)
:- support for wire bonding
:- support for module testing
:- glue dispensing for butt-joint and for ceramic inserts
:- alignment of two half-modules
:- support during room temperature curing of the joint
After some discussions and meetings at MPP and HLL:
 Ladder Assembly Sequence
 Concept of the “Ladder Universal Jig”, drawn by Tscharlie

## Slide 6: Step 0: preparation
:- Wafers are cut, modules go through final clean at HLL
:- pack modules into 4’ gel-paks, send to IZM for FC
:- FC on Si support jigs (designed and made by IZM)
:- FC and X-ray at IZM
:- in gel-paks back to HLL  inspect, X-ray on sub-sample
:- place into transport jig, label ..
 From this point, the module stays in the jig!
Base module jig
base module jig
transport base
cover

## Slide 7: Step 1: SMD
:- modules in transport box shipped to NTC for SMD
:- disassemble, remove base module jig transport box
:- Vacuum channels in base module jig fit to vacuum channels at NTC equip.

## Slide 8: Step 2: Probe station
:- cooling plate
:- optional cooling during test, sensor held by vacuum
:- cover over sens. area

## Slide 9: Step 3: Kapton attachment
:- attach light weight kapton jig to base module jig
:- all necessary features for kapton vapour phase soldering and wire bonding

## Slide 10: Step 4: testing and storage

## Slide 11: Step 5: the tricky part – turn it around..

## Slide 12: Step 5: the tricky part – turn it around..

## Slide 13: Step 5: the tricky part – turn it around..

## Slide 14: Step 5: the tricky part – turn it around..

## Slide 15: Step 5: the tricky part – turn it around..

## Slide 16: Step 6: screen printing of the adhesive – opt. 1

## Slide 17: Step 6: screen printing of the adhesive – opt. 2

## Slide 18: Step 7: alignment and gluing

## Slide 19: Step 8: turn back

## Slide 20: Step 8: turn back

## Slide 21: Step 8: turn back

## Slide 22: Step 9: test

## Slide 23: Step 10: pack and store

## Slide 24: In Summary
:- please allow me not to summarize … puuuuuu
:- this is a concept, a very elaborated one, but still a concept
:- we have start with prototypes, practice (screen printing, handling ..), and optimize