## Slide 1: PA and Origami Status and Plan
Koji Hara (KEK)
Jan. 21, 2015
(updated on Jan. 23)
7th Belle II VXD Workshop and 18th International Workshop on DEPFET Detectors and Applications
Comments added after the session on Jan. 21
  Additional information about reflow at REPIC
  P. 6-7, Appendix P. 18-19
  Comment on the PA0+Origami reproduction schedule
  P. 11
  Batch0 with 1-layer
  P. 12

## Slide 2: PA0 and Origami
PA0 problems
  Cracks
  Shorts between pads
  Design to keep
    bonding pad length
    Available size of the bias hole
  Shrinkage (likely by heat)
    0.1-0.2%
    Not uniform??
Origami status
  Produced at Taiyo
  Preserved as sheets before cutting and electrical test
  Need to confirm Origami+z prebending method (performed by SVD group)
    Any modification necessary from the previous method?
Electrical test of PA0
  4 probe measurement possible for PA0 alone and PA0 glued on Origami
  Test condition to suppress probe traces
Practice whole procedure
PA0 from Taiyo  Electrical test @ Daiei  PA0 + Origami gluing @ Taiyo
 Electrical test @ Daiei  assembly@REPIC  test by APVDAQ@KEK

## Slide 3
Crack
  Long straight design
  Acrylic glue
  Laser cutting
Shorts
Still
36KΩ!
Carbonized glue under the top base by laser cut touches the metal line
New design
covered by coverlay on the bottom layer

## Slide 4: PA0 latest Design by Taiyo
Released for review by SVD group on Jan. 19
  Followed our request on the pad lengths and bias hole positions
dxf files on PATaskForce twiki page

## Slide 5: Discussion item: Shrinkage of PA0 length
Measurement of length between alignment marks on PA0
  Design : 124.3 mm
  At HEPHY
    Assembled by REPIC : 124.07 – 124.17 mm
  At IPMU
    Mockup (full process but no reflow soldering) : ~124.22-25 mm
    Assembled by REPIC : ~124.12-17 mm
Must understand the shrinkage
Need to change the process?
Non-uniform shrinkage ?
“No problem” by HEPHY reflow (Vapor Phase)
  No or uniform shrink?
Corresponding to the typical thermal shrink rate of polyimide
～0.1 % @ 200 ℃
  Seems more or less inevitable

## Slide 6: Heating in Origami+PA0 production
Procedure
  1. Gluing of PA0 coverlays and 2 layers by heat pressing at Taiyo
  2. Glue PA0 and Origami by heat pressing at Taiyo
  3. Reflow soldering of Origami at REPIC
  Q) Is it possible to do 1  3  2 ?
  A) No. Heat pressing becomes impossible after mounting parts on Origami. (Taiyo)
Heat pressing at Taiyo
  Gluing of coverlays, top + bottom, PA0 + Origami
  160 degrees, 40 minutes, 45-50kg/cm2
  The tolerance of the positions of the alignment marks and the bonding pads (>100µm) are general tolerance of 0.2mm if no special request given
Reflow soldering at REPIC (Saiden Device)
  220-225 degrees (maximum temp.), 4 minutes
  2 times for top and bottom surfaces
Q) Is it possible to change the method to VPS?
A) Checking the possibility. (Saiden Device does not have VPS machine.)

## Slide 7: Reflow at REPIC (Saiden Device)
Origami fixed on G-10 board with Kapton tapes
Board carried by conveyor
Send to oven
  Heated by hot air
  3 heat cavities set to
    160 degrees (for preheat)
    160 degrees (for heating up)
    220-225 degrees (for soldering)
  Temperature monitored by sensors inside
  Temperature profile data in appendix
  20-30 sec. with temp. > 200 degrees
  Length : about 3m
  Speed 3m/4min.
～3m

## Slide 8: Electrical Test
Verified possibility to test PA0 alone and PA0 after glued on Origami
  4 probe resistance measurement
  With “soft landing” probe
Probe traces and scratches found  need to be solved
  Visited the company again and discussed, tested PA0 pieces together on Dec. 19
  Scratches not observed on Tokai Denshi PA
    Tested as a sheet  seems to be better and uniform vacuum chucking
    Different machine (No. 2 and No. 4)
  Some areas have clear traces, some areas do not
    Not depending on where the PA0 placed on the stage
     Dependence on the machine or the probe?
  Probe speed can be changed
  Improvement seen by reducing the probe speed
Testing PA0 test with different speeds
  Machine to be used for the final test condition
    Isolation >100MΩ@10V, connectivity < 100Ω

## Slide 9
Probe speed 30%
Speed 10%
Testing more slower speed
  with the same machine used for Tokai Denshi PA test
    recording the S/N and use count of the probes
  Down to speed setting of 3～5%
  Tested PA0 will be returned next week  will be checked by microscope
Still need a few more iteration to reach to the final test condition
  PA0 alone and PA0 + Origami
  could be a bottle neck, if failed to converge to the good test condition
  1-2 months

## Slide 10
 | Taiyo | Saiden | Hayashi Watchworks | Hayashi Watchworks
 | Origam FPC | Reflow soldering | APV gluing + wire bonding | H/C-shape gluing
weeks |  |  |  | 
1 | Batch0 12 pcs |  |  | 
2 |  |  |  | 
3 |  |  |  | 
4 |  |  |  | 
5 |  | 12 pcs |  | 
6 |  |  |  | 
7 |  |  |  | 
8 |  |  | 12 pcs | 12 pcs
9 |  |  |  | 
10 |  |  |  | 
1 | Reproduction 141 pcs |  |  | 
2 |  |  |  | 
3 |  |  |  | 
4 |  |  |  | 
5 |  | Batch1 47 pcs |  | 
6 |  |  |  | 
7 |  |  | 47 pcs | 47 pcs
8 |  |  |  | 
9 |  | Batch2 47 pcs |  | 
10 |  |  |  | 
11 |  |  | 47 pcs | 47 pcs
12 |  |  |  | 
13 |  | Batch3 47 pcs |  | 
14 |  |  |  | 
15 |  |  | 47 pcs | 47pcs
16 |  |  |  | 
17 |  |  |  | 
18 |  |  |  | 
REPIC
Taiyo
Origami+PA0 production + assembly (incomplete) schedule
H/C-shape gluing work at REPIC is not yet agreed / decided.
  REPIC proposes to do by them before APV gluing/bonding for safety of connectors
  Has been done at KEK by Tsuboyama-san and Sasaki-san after delivery from REPIC
Electrical test and Origami +Z bending are not included yet in the schedule
batch0
Mass reproduction
-z 6
ce 4
+z 2
1 batch =
-z 16
ce 22
-z 9

## Slide 11: Input of Discussion : Taiyo PA0 + Origami Possible Schedule
PA0 new design evaluation 1-2 weeks
PA0 test production 2 weeks
Evaluation of PA0 sample 2 weeks
  Crack-free
  no shorts
  Bonding pad lengths, bias hole effective size
  Stress test
Sign-off the final specification of PA0 + Origami reproduction 1-2 weeks
accompanied by
  Electrical test condition with minimum probe traces
  Origami +z bending method
  H/C-shape gluing specification update
Origami+PA0 batch0 production 3 months
Evaluation of batch0 Origami 1 months
Mass reproduction (3 batches) 3-5 months
2015 Jan
2015 Feb
2015 Mar
2015 June
2015 July
2015 Oct-Dec.
contingency of a few weeks
Assumed electrical test 2 weeks
Produce more PA0 in the preproduction to use them for batch0 (if successfully)
Ask to try to reduce duration of rigami+PA0 production in batch0
In addition, Batch 0 production with 1 layer PA0 is preferable

## Slide 12: 1-layer 3-row PA0
Production drawing prepared by Tokai Denshi for SVD group’s review
  Distributed to PA taskforce Jan. 16
  Uploaded to PATaskForce twiki page
Planning test production (10 pieces)
  1.5 month for production
  Payable by KEK budget in 2014 JFY
    If delivery before end of Mar.
     If approval of SVD made by mid of Feb. at latest.
Open issue : how and who glue it on Origami
Batch0 production with 1-layer 3-row PA0 is preferable

## Slide 14: PA flex
Signoff the full specification document by SVD
  Fix Pad width minimum limit by SVD measurement
    30 µm?
    By CMM at IPMU
  Bonding on PA glued on Airex + Origami @ IPMU
   result by Friday?
Must agree with Tokai Denshi on the full specification document translated in Japanese
  Now translating
By end of Jan.

## Slide 15
Type | #required | 11-Mar | 17-Mar | 24-Mar | 30-Mar |  | 3-Apr | 7-Apr | 10-Apr | 14-Apr |  | #delivery
PA0 | 10 | 10 |  |  |  |  |  |  |  |  |  | 10
3PF1 | 30 | 10 | 10 |  |  |  |  |  | 10 |  |  | 30
3PF2 | 30 | 10 | 10 |  |  |  |  |  | 10 |  |  | 30
3PB1 | 30 | 10 | 10 |  |  |  |  |  | 10 |  |  | 30
3PB2 | 30 | 10 | 10 |  |  |  |  |  | 10 |  |  | 30
PA1 | 200 |  | 40 | 20 | 20 |  | 20 | 20 | 40 | 40 |  | 200
PA2 | 200 |  | 40 | 20 | 20 |  | 20 | 20 | 40 | 40 |  | 200
PF1 | 100 | 20 |  | 20 | 20 |  | 20 | 20 |  |  |  | 100
PF2 | 100 | 20 |  | 20 | 20 |  | 20 | 20 |  |  |  | 100
PB1 | 100 | 20 |  | 20 | 20 |  | 20 | 20 |  |  |  | 100
PB2 | 100 | 20 |  | 20 | 20 |  | 20 | 20 |  |  |  | 100
 |  |  |  |  |  |  |  |  |  |  |  | 930
Assumed the production start on Feb. 2
Not final in terms of payment and contract
# pieces in batch may change
  To wait for budget open in the beginning of JFY
  To pay more with 2014 budget
    KEK and support from IPMU
  Still discussing
Need to be delivered in May to wait for the budget open in the beg. of 2015 fiscal year.
PA delivery schedule
ver. 20150109 by Tokai Denshi

## Slide 17
 |  |  | 2014 |  |  |  |  | 2015 |  |  |  |  | 
 | 品名 | 要求数量 | 3月11日 | 3月17日 | 3月24日 | 3月30日 |  | 5月11日 |  |  |  |  | 納入数量
 | PA0 | 10 |  |  |  | 10 |  |  |  |  |  |  | 10
 | 3PF1 | 30 |  |  |  |  |  | 30 |  |  |  |  | 30
 | 3PF2 | 30 |  |  |  |  |  | 30 |  |  |  |  | 30
 | 3PB1 | 30 |  |  |  |  |  | 30 |  |  |  |  | 30
 | 3PB2 | 30 |  |  |  |  |  | 30 |  |  |  |  | 30
 | PA1 | 200 | 50 | 50 | 50 | 30 |  | 20 |  |  |  |  | 200
 | PA2 | 200 | 50 | 50 | 50 | 30 |  | 20 |  |  |  |  | 200
 | PF1 | 100 | 30 | 30 | 40 |  |  |  |  |  |  |  | 100
 | PF2 | 100 | 30 | 30 | 40 |  |  |  |  |  |  |  | 100
 | PB1 | 100 | 30 | 30 | 40 |  |  |  |  |  |  |  | 100
 | PB2 | 100 | 30 | 30 | 40 |  |  |  |  |  |  |  | 100
 |  |  |  |  |  |  |  |  |  |  |  |  | 930
PA delivery schedule assuming the start on Jan. 26
Better for budget
  Consume maximum in 2014 JFY budget
Making final signoff and agreement with the full document seems not realistic

## Slide 18: Appendix added on Jan. 22: Temperature profile data of the reflow by REPIC (Saiden Device)
Measured by sensors put on a FR4 board
  Board size ～ A4
  6 sensors equipped on the board
  Recorded the temperature by time in the reflow oven
(Which sensor is at where is not given)