## Slide 1: VXD Radiation and Environmental Monitoring
L. Bosisio, P.Cristaudo, L.Lanceri, L.Vitale
INFN and Univ. Trieste
+ Electronics&Detectors group at Elettra, Trieste

## Slide 2: Outline
Radiation monitor & beam abort
  scCVD diamond sensors, electronics
  Background simulations
  Beast Plans
Temperature monitoring
  Optical Fibers Sensors (FOS)
  NTC thermistors
Humidity monitoring/interlock
SVD+PXD interlocks: PLC
Parallel session on Friday 10.30: discussion
Still ongoing in Trieste: sensor testing - not covered in this talk

## Slide 3: Radiation monitoring & Beam abort

## Slide 4: scCVD radiation sensors
6 + 6 sensors
on SVD support cone
close to L3 rings
4 + 4 sensors
PXD-beam pipe
Shielded
diamond
sensors
3+15 m (3+40 m) cables
Voltage sources (150÷500 V)
picoAmmeters

## Slide 6: Mechanics, cabling
Compact packages:
  1) SVD 12 x 20 x 3.1 mm3
  multi-layer package, Rogers laminate ok for SVD (we already have 28 pc)
  2) PXD 12 x 10 x 3.3 mm3
  Same material, with perpendicular or parallel mounting (simulation)
Cabling (~20 m)
  3 m coaxial thin (HS SM47LSFH) from sensors to DOCKs
  + 15 m (or more) HS S_04162-B60, double shield

## Slide 7: Expected radiation from simulations
VXD radiation/abort sensors included in simulation geometry, thanks to Martin Ritter MPI (Nov. 2014). Now Gianluca Inguglia DESY is taking over.
Particle fluence (e-)
Energy deposit
Preliminary
Preliminary

## Slide 8: Location for “beam pipe” sensors: “perpendicular” option

## Slide 9: Modified package for “perpendicular”

## Slide 10: First tests of “final” package

## Slide 11: After MIPs tests: dosimetry
Sr90 point-like source, 3 MBq, at different distances, compared to FLUKA simulations
d [mm]

## Slide 12: Long term stability studies
1 day
HV = 50-200 V

## Slide 13: Current (dose rate) vs HV
unexpected voltage dependence (no saturation?)
on both tested sensors (Micron)
at all distances
cross checked with different picoammeters
I(nA) vs HV(V)

## Slide 14: Diamond sensors are not that simple
Simplest model: just “ionization chambers”, but…
pCVD, known solid-state effects:
  “pumping” or “priming” (filling of long-lived “deep” traps)
  on/off transients (“shallow” traps, space charge and polarization)
  electrodes: complicated metal-diamond interfaces
    blocking vs. ohmic contacts, depending on materials & preparation
    charge injection and “photoconductive gain >1” are possible
scCVD: less defects, more stable behavior
  However similar (smaller) effects are possible
  Seen in our lab, to some extent (transients, gain?)

## Slide 15: Sensors: plans
One “final” Cividec prototype ordered recently
Final acceptance tests will include:
  I-V without irradiation, both polarities
  Charge collection efficiency, MIPs (90Sr β-source)
  Currents/dose rates measurements (90Sr β-source)
  Measure electric field uniformity inside with TCT (Transient Current Technique) + 241Am α-source
At the moment Cividec
  seems to be the most expert commercial provider of complete, metallized (and preliminarily tested) sensors
Other solutions are still open (discussion Friday)

## Slide 16: Cividec acceptance tests
←1nA
←10pA
500V
10ns
→ ←
β-source 54MBq
Distance ~1cm
@ 200V
Drift time
α-source
3600 s

## Slide 17: Progress in electronics design
Collaboration with electronics&detectors group at Elettra, Trieste
  G. Cautero, D.Giuressi, R.H.Menk + F.Vulpone (student)
Modular design, FPGA + Memory + ADCs + HV + Ethernet
  project’s kernel: ALTERA DE3 evaluation board, with a STRATIX III FPGA
  External memory: Transcend 1GB DDR2 RAM
  3 peripheral boards will be connected with DE3
    board with 4 ADCs, HSMC connector
    board with 2 Ether-W-ease (Ethernet), 2 GPIO connectors
    board with 4 DACs and 4 HV diamond bias outputs
Schedule plans:
  “Final” prototype available for test with 4 sensors at BEAST phase 1
  “Production” modules ready for BEAST phase 2

## Slide 18
BLM Digitizer Card (“logger”):
ADCs, circular buffers, running sums
BLM Abort Card:
masks, majority logic,
thresholds (4 levels)
Desired functionality:
similar to Fermilab BLM
Analog front end + ADC
derived from
Elettra AH510 Pico ammeter

## Slide 19: Rad.Mon+Abort, 4-channel Box
3 Sums
4 Abort Levels
Circular
Buffers
Ethernet I/O
Ext. Clock (100 kHz)
Abort from SuperKEKB
Timing buffer
Timing
Control
Abort
HV
SuperKEKB status
Thresholds
etc.
(for each
input channel)
From 4 diamond sensors
to 4 sensors
to SuperKEKB
ADC
FPGA

## Slide 20: Design and implementation plans
DE3/FPGA
Ether-W-ease
Ether-W-ease
ADCs
DACs
Input currents
HV outputs
Ethernet I/O
Read and write the external RAM memory using the FPGA
Design and prototypes of the ADC’s and DAC’s boards
Design and prototype board with two Ether-W-ease
VerilogHDL program for the FPGA
  Synchronization, buffers, running sums, thresholds, abort logics
SuperKEKB timing
Abort signals

## Slide 21: Environmental monitoring

## Slide 22: SVD FOS fibers system (baseline)
Layer 6: 16 ladders
16 fibers x 4 (+1) FBGs
DOCKs
(4x1)couplers
…
…
Layer 5: 12 ladders
12 fibers x 3 (+1) FBGs
…
…
Layer 4: 10 ladders
10 fibers x 2 (+1) FBGs
…
…
Layer 6
4 couplers
Layer 5
3 couplers
Layer 4
3 couplers
Interrogator
(16 inputs)
…

## Slide 23: SVD FOS fibers system -2 FBGs/origami
Layer 6: 16 ladders
16 fibers x 7 (+1) FBGs
DOCKs
(4x1)couplers
…
…
Layer 5: 12 ladders
12 fibers x 5 (+1) FBGs
…
…
Layer 4: 10 ladders
10 fibers x 3 (+1) FBGs
…
…
Layer 6
4 couplers
Layer 5
3 couplers
Layer 4
3 couplers
Interrogator
(16 inputs)
…

## Slide 24: FOS temperature sensors: status
Recent purchases (delivered last week and yesterday):
  16-channel interrogator, Micron Optics (Pisa)
  Prototypes with different numbers of sensors per fiber, from 2 providers:
    Micron Optics
    Smart Fibres
  Final purchases
    38 fibers
    Early 2015
Environmental chamber for calibrations
  [−30°C ÷ +70°C], available in Trieste; help from Santander welcome!

## Slide 25
SM 225-800 16-channel interrogator, Micron Optics
(4x1)coupler
(2x1)coupler
Test fiber received from Santander
Fiber with 8 sensors
Fiber with 5 sensors

## Slide 26: NTC ELMB readout, preliminary design
sensors
docks
Temperature:
up to
6 x 16 = 96
NTC thermistors
Box in Electronics Hut
PC
CANbus
ELMB
ELMB
ELMB
Power supply
8 twisted pairs
(no shield)
Shielded
twisted
pairs (16)
(shield
grounded
at docks)
Current
sources
+ interlocks
2 interlock signals
6 x 2 signals
to VXD
Interlock
logics
EPICS
Slow Control
≈ 3÷4 m
8 twisted pairs
(no shield)
Small
connect.
board
connectors
+ jumpers

## Slide 27: NTC: full prototype chain tested
Motherboard
ELMB
board
CANbus
adapter
Current sources
Interlock comparators
100k NTC thermistors
- Betatherm (TS)
- Murata (Vienna)

## Slide 28: Interface, software and stability
New Kvaser CANbus-USB interface installed
Software
  Stand-alone debugging program
  LabView + CANopenOPCServer 2.9.7.4
  EPICS: expert visit welcome
Uniformity and stability tests on 11 thermistors
  OK !

## Slide 29: Final motherboards design, completed
sensors
docks
Temperature:
up to
3 x 32 = 96
NTC thermistors
Box in Electronics Hut
PC
CANbus
ELMB
ELMB
Power supply
4 x 8 twisted pairs
(no shield)
twisted
pairs (32)
Current
sources
+ interlocks
4 interlock signals
3 x 4 signals
to VXD
Interlock PLC
EPICS
Slow Control
≈ 3÷4 m
3 Small
connect.
boards
connectors
+ jumpers
ELMB
Current
sources
+ interlocks
Current
sources
+ interlocks

## Slide 30: “Final” NTC ELMB design readout and interlock board
Input
32 NTC
interlock
to PLC
ELMB
Current sources and interlock comparators
CANbus connector
Reference resistors

## Slide 31: Humidity monitoring & interlock

## Slide 32: Dew Point Sensors (interlock @ -30°C)
Vaisala DMT242B
Dew Point Transmitters
[-60, +60]oC dew point range
Edgetech Dewmaster
Chilled Mirror Hygrometer
(for calibrations)
Rotameter
flux meters
Pump
Sniffing pipes
Main components
received
Prototype planned
(1st quarter 2015)

## Slide 33: SVD + PXD Interlocks
?
?
?
Temperature (NTCs)
Dew point (Vaisala)
Sniffer Pump(s)
Cooling plant failure
Beam Abort
Fire alarm?
…etc
PXD power supplies
VXD power supplies
(not segmented)
…etc
Programmable Logic Controller (PLC)
Schneider M340 – cpu BMX P34 2030
Modbus/EPICS drivers exist! expandable
kit (p.s. + cpu + I/O + software) purchased
Local expertise: INFN technician
INPUTS
OUTPUTS
INTERLOCK
LOGICS

## Slide 34: Summary
Good progress on:
  Radiation monitor & beam abort
  Temperature monitoring
  Humidity monitoring/interlock
  SVD+PXD interlocks: PLC
Documentation to be prepared!!!
Parallel session on Friday 10.30: discussion
A very busy period also for us …

## Slide 35: Approximate schedule
Temperature & humidity | 2014 |  |  |  | 2015 |  |  |  | 2016 | 
What | 1 | 2 | 3 | 4 | 1 | 2 | 3 | 4 | 1 | 2
T&H Specifications, sign up |  |  |  |  |  |  |  |  |  | 
Fibers, mechanical layout |  |  |  |  |  |  |  |  |  | 
Fibers, assembly tests (SVD) |  |  |  |  |  |  |  |  |  | 
Fibers, multiplexing scheme |  |  |  |  |  |  |  |  |  | 
Fibers design sign up |  |  |  |  |  |  |  |  |  | 
Procurement, assembly |  |  |  |  |  |  |  |  |  | 
installation (PXD & SVD) |  |  |  |  |  |  |  |  |  | 
Thermistors, lab tests |  |  |  |  |  |  |  |  |  | 
Mechanics & cabling design |  |  |  |  |  |  |  |  |  | 
Procurement |  |  |  |  |  |  |  |  |  | 
assembly, installation |  |  |  |  |  |  |  |  |  | 
Radiation | 2014 |  |  |  | 2015 |  |  |  | 2016 | 
What | 1 | 2 | 3 | 4 | 1 | 2 | 3 | 4 | 1 | 2
Specifications, sign up |  |  |  |  |  |  |  |  |  | 
Simulations |  |  |  |  |  |  |  |  |  | 
Dosimetry, characterization |  |  |  |  |  |  |  |  |  | 
Diamond sensors choice |  |  |  |  |  |  |  |  |  | 
Mechanics & cabling design |  |  |  |  |  |  |  |  |  | 
Electronics specifications |  |  |  |  |  |  |  |  |  | 
Electronics design & proto. |  |  |  |  |  |  |  |  |  | 
Overall design, sign up |  |  |  |  |  |  |  |  |  | 
Components, procurement |  |  |  |  |  |  |  |  |  | 
"final" prototypes (4 ch.) |  |  |  |  |  |  |  |  |  | 
BEAST, 1 preliminary tests |  |  |  |  |  |  |  |  |  | 
BEAST, 2 complete test (4 ch) |  |  |  |  |  |  |  |  |  | 
electronics production & test |  |  |  |  |  |  |  |  |  | 
sensors installation & cabling |  |  |  |  |  |  |  |  |  | 
electronics installation |  |  |  |  |  |  |  |  |  | 

## Slide 36: Back-up slides

## Slide 37: Diamond sensor characterization
Label | Provider | sc/p CVD | Geometry mm^3 | pads | Validation 90Sr beta
PD | DDL | sc | 4.7x4.7x0.15 | 2x2 | OK
DC1 | Cividec | sc | 4.7x4.7x0.50 | 3x3 | OK
DC2 | Cividec | p | 5.0x5.0x0.50 | 3x3 | TBD
DM1* | Micron | p | 5.0x5.0x0.50 | 1 | OK
DM2* | Micron | sc | 4.7x4.7x0.50 | 1 round | OK
DM3 | Micron | p | 5.0x5.0x0.50 | 1 | Rejected
DM4 | Micron | sc | 4.7x4.7x0.50 | 1 | OK
DM5 | Micron | sc | 4.7x4.7x0.50 | 1 | OK
DM6 | Micron | p | 5.0x5.0x0.50 | 1 | TBD

## Slide 38: Diamond sensors
PD
DM1
DC1
DM4

## Slide 39: Test tools
Static characterization of the samples (I-V and C-V)
Single channel Amptek chain: A250 preamp. (2.5 ns risetime) + PX5 digital shaper (semi-gaussian equivalent time 40 ns – 40 µs).
A pointlike 90Sr source (~2 MBq), beta spectrum hardened by magnetic field and collimators
Tests repeated at DESY 2-6 GeV electron beam

## Slide 40: MIP Characterization of scCVD sensors
Single electrons, 1-2 MeV (MIP): Landau peak vs HV
MPV vs HV

## Slide 41: Belle II-VXD Radiation Monitoring and Beam Abort
Main requirements
  Measurement of instantaneous dose rates & integrated doses
    Sensitivity ≈ 1 mrad/s = 10 μGy/s, sampling rate ≈ 100 KHz
  Beam Abort for excessive beam losses affecting PXD, SVD
    Programmable thresholds, depending on accelerator conditions (filling injection, stable beams, machine development, …)
    “Fast” (10 μs = 1 accelerator revolution) Beam Abort trigger
    “slower” Beam Abort triggers, based on digitally filtered signals (averages over programmable numbers of samplings)
Conceptual design
  Based on experience from Belle, BaBar, CDF, LHC
  scCVD diamond sensors, measurement of currents:
    Reference: typical pCVD sensor (10 × 10 × 0.5) mm3 @ 500V: 1 nA = 7 mrad/s = 70 μGy/s
    Noise should be limited to a few pA, in current measurements
  4 + 4 sensors located near PXD (very limited space)
  6 + 6 sensors located near SVD
  Background dominated by electrons

## Slide 42: VXD Monitoring & Beam Abort electronics
5 modular Boxes, dealing with 4 scCVD sensors each:
  4 individual HV bias supplies
  Currents digitization with 4 (16-bit) ADCs at 160 MHz, with oversampling
  Data averaging and buffering for
    Continuous Monitoring (averaged data, 1-10 Hz read-out)
    Post-mortem (abort) analysis (full data, frozen circular buffers)
  Running Sums, Programmable Thresholds, Majority logics for Abort signals
  Ethernet interfaces (2) for programming and data read-out
Compatibility with the SuperKEKB Abort System requires:
  Sampling synchronized to SuperKEKB revolution period (10 μs)
  “SuperKEKB status” register, set via EPICS (for the selection of thresholds)
  2 final output “fast abort” signals, correlated with LER and HER
  600 ms deep data buffers, read out upon an abort, internal or external
  Synchronization with SuperKEKB aborts: digitization of 2 beam currents, and “abort confirmation” digital signal from SuperKEKB

## Slide 43: Similar approach: BLM at LHC
FPGAs
Running Sums
Thresholds

## Slide 44: NTC current sources and interlock signals
6 small boards, each:
  Current sources for 16 NTCs
  16 comparators with hardwired thresholds
  2 interlock signal outputs (each: OR of 8 channels)
Built-in redundancy
  2 NTCs per measurement point
  Each contributes to a separate interlock OR
Design completed, prototype channel tested

## Slide 45: ELMB-based NTC readout
In one box, with 3 x 32 NTC channels modularity:
  1 power supply board, 3 motherboards with ELMB and CANbus adapter, 3 x 2 current source (16-channel) boards, 3 x 2 x 16-channel connectors