## Slide 1
VXD Cooling Requirements
IBBelle Overview
Commissioning at CERN (ATLAS version)
Software Control
Ordering of Parts
Location of IBBelle at KEK
Schedule and Responsibilities
Status of IBBelle

## Slide 2
VXD Cooling Requirements & Specs
Minimize material in acceptance region
SVD requirements:
coolant temperature for APV25 @ -20°C for SNR improvement
total dissipated power ~ 700 W
PXD requirements:
sensor temp. < 25°C, ASICs < 50°C : cooling block @ <-20°C
dissipated power ~ 360 W
Cooling power ~ 1.6 kW at VXD (incl. ~ 540 W parasitic heat influx)
have to add heat loss through 15 m of transfer lines → ~ 2-3 kW
Dry Volume (nitrogen):
dew point inside VXD volume: ~ -30 °C
Outside of VXD: room temperature ( ~ 23°C) at the CDC
may need a thermal enclosure towards CDC, well isolated

## Slide 3
SuperKEKB/Belle II Schedule
JFY2018
2018
2014
JFY2014
2015
2016
2017
JFY2015
JFY2016
JFY2017
Japan FY
now
Calendar
Power restriction
in summer
phase 1
phase 2
phase 3
Start-up
QCS, Belle II install
VXD install
Current plan ongoing
w/o QCS
w/o Belle II
w/ QCS
w/ Belle II (no VXD)
w/ full Belle II
phase 1
phase 2
phase 3
SuperKEKB Construction
Start-up
QCS, Belle II
VXD install
Original plan
w/o QCS
w/o Belle II
w/ QCS
w/ Belle II (no VXD)
w/ full Belle II
phase 1
phase 2
Start-up
QCS, Belle II install
VXD install
w/o QCS
w/o Belle II
w/ QCS
w/ Belle II (no VXD)
JFY2014 Operation budget cut by 12 Oku-Yen
(A) JFY2015 budget 60 Oku-Yen
SuperKEKB Construction
SuperKEKB Construction
Power restriction
in summer
Power restriction
in summer
phase 3
w/ full Belle II
Power restriction
in summer
CO2 Cooling is needed in Tsukuba hall already for BEAST 2
“conceivable” schedule
shown during B2GM in Nov. 2014
“C4”
JFY2015 Operation
(small) budget cut

## Slide 4
DCS:
TTa24 - TTn24
(FTa24 - FTn24)
EH122
TS122
MVa24 - MVn24
04
06
FL018
⅜”
AV017
20
36
BD016
PT116 / PT316
TT116 / TT316
MV018
MV036
14 staves (a-g),(7 flow pairs)
(10x BW flow / 4x FW flow)
Junction box
Two-stage dry volume
PXD / SVD
Transfer lines
(~ 100 m)
HX036
½”
⅜”
Dummy load (testing only)
BD020
PT120 / PT320
TT120 / TT320
TTz20 (DCS)
BD036
PT136 / PT336
TT136 / TT336
TTz36 (DCS)
MV035
EH117
TT117
TS117
PT510
VP512
VP510
EV508
nc
HX012
16
FL017
MV017
Manifold box
12
10
12
10
08
PT512
PT610
VP612
VP610
PT612
TT510
TT610
nc
EV608
Cables and inter connection pipes
DN40 vacuum
Vacuum system
DN40
¼”
⅜”
¼”
½”
CO2 System Overview
Vacuum isolated
vacuum isol.
CO2 Plant „IBBelle“
(IBL + Belle)
Flex lines
~15 m

## Slide 5
AC042
LP101
vent
evacuate
6
8
FT106
FL104
⅜”
EH106
TT106
TS106
EH101 / EH102 / EH103
TT101 / TT102 / TT103
TS101 / TS102 / TS103
PT101 / PT102 / PT103
HX150
System A
LT142
LT342
FT306
FL304
⅜”
FL306
VP056
50
4
40
12
44
44
46
48
PV110
BD150/ PT150/ TT150
¼”
PT104
TT104
BD108
PT108
TT108
CO2 from experiment
CO2 to experiment
42
PT142
EH142/143
TT142/143
TS142/143
EH342/143
TT342/143
TS342/143
PV108
PV144
HX148
TT148
SV042
SV043
MV042
FL144
MV041
TT146
AV108
Freon chiller A 200
System B
10
LP101
EH301 / EH302 / EH303
TT301 / TT302 / TT303
TS301 / TS302 / TS303
PT301 / PT302 / PT303
4
FL344
BD304
PT304
TT304
MV306
6
8
EH306
TT306
TS306
BD308
PT308
TT308
AV308
PV308
PV310
PV344
46
TT346
HX350
HX348
LP301
Fill port
nc
nc
nc
nc
nc
nc
MV050
MV054
MV052
MV056
BD050
PT050
EV148
EV348
nc
nc
50
BD350/ PT350/ TT350
SV040
MV040
SV041
BD112
10
MV058
NV110
MV110
MV310
nc
CV142
nc
CV342
nc
nc
nc
nc
nc
nc
nc
FL106
Fill port
MV106
Cold CO2 line
Cold R404a line
Warm service line
(Cold lines require 32mm insulation)
no
NV310
no
¼”
¼”
½”
½”
½”
½”
½”
bs
CV149
nc
50
PT149
TT149
48
TT148
nc
50
CV349
PT349
TT349
Freon chiller B 400
MV043
PT342
IBBelle

## Slide 6
Installation at CERN (ATLAS Experiment): the “twin” of IBBelle
Accum.
rack
IBBelle at CERN
Construction of IBBelle (ATLAS) is finished, commissioning with dummy heat load is done
Some problems getting down to -40°C
(not important for the Belle version
IBL now cooled down

## Slide 7
Mechanics: Production drawings mostly up-to-date
Parts list largely established, still visit at CERN necessary
MPI working on this: Sven Vogt
Swagelok part list done, ordering ongoing, first parts have arrived
Control (PCLs etc): mostly documented.
MPI working on this: Reinhard Sedlmeyer
Further visit to CERN necessary
Parts being ordered through CERN, parts are coming in
Ordering of Parts for the CO2 Plant

## Slide 8
MLI (super) insulation
1.6 mm OD /
1.0 mm ID concentric inlet tube
4 mm OD / 3 mm ID outlet tube
18 mm OD corrugated vacuum tube
~ 12m long
CO2 Vacuum Flex Lines
MLI
admissible bending radius:
~ 50 mm
CERN crew offers help building the tubes
Flex line design (by CERN/NIKHEF)

## Slide 9
Construction and Commissioning
IBBelle for VXD will be an almost EXACT copy of IBBelle (ATLAS)
Variants possible in the transfer system from IBBelle to VXD
(include results from Thermal Mockup)
Up-to-date 3D CAD model and production drawings exist
Ordering of parts according to lists established at MPI (based on material from CERN/Nikhef for ATLAS IBL) by MPI
Construction at MPI starts in 2015 (workshop is prepared), including pressure tests, control electronics, etc.
Commissioning will be done with the help of CERN/Nikhef crew (exact dates to be organized) by the end of 2015 at MPI, including engineers from DESY, KEK and SVD (tbd)
Certification by TÜV (process being prepared at present)

## Slide 10
PVSS -> EPICS
Remark: This is a desired option, not a necessity

## Slide 11
PVSS GUI for IBBelle
PVSS sets operation points (T, CO2 flux) and monitors the operation
Supervising layer (+GUI) will be
re-written in
EPICS / CSS
(Michael, Bogdan, Thorsten)
All critical control / operation functions are formulated in PLC and will not be touched
Plant A

## Slide 12
IBBelle Form Factor: ATLAS Model
40’ Container, to be placed outside of Tsukuba hall
Advantage: unit can be assembled and commissioned at MPI, no need to take it apart for transport, no reassembly
Requirements at KEK: Need earth-quake-safe hut around IBBelle, need water/electricity outside hall

## Slide 13
Services for IBBelle
European standard P/S at KEK, needs to be reactivated
Requirement: 400 [V] ± 10 % (filtering of V spikes)
Water for the chiller cooling: 20 l/min, room temp OK, only filter needed, connections to be specified
N2: 50 l/min (Gas), room temp OK
Air conditioning: room temp ± 3°C, rel. humidity < 70 %
Transfer lines & connection standards still need to be specified (depending on location of IBBelle)

## Slide 14
Phase 2 / PXD Integration
VXD Assembly clean room in B1

## Slide 15
VXD Clean Room in B1
Tscharlie Ackermann
Should already be available for BEAST Phase 2 assembly

## Slide 16
Proposal for a “Small” Hut

## Slide 17
Proposal for a somewhat larger Hut

## Slide 18
Proposal for a somewhat larger Hut
Possible location of IBBelle: behind Tsukuba Hall
Transferline: L_max = 112 m (to B1), 80 m (to Belle)

## Slide 19
Schedule for the CO2 Plant
IBBelle at KEK in summer of 2016

## Slide 20
Responsibilities, Operation at KEK
Lead institute (main responsibility): MPI Munich
Collaborating institutes: DESY + SVD (tbd)
Installation at KEK: MPI, DESY, SVD (= “VXD”)
Commissioning: VXD + KEK
Agreement with KEK (Nov. 13 B2GM meeting):
KEK will operate and monitor IBBelle
standard service tasks executed by KEK
(e.g. filter cleaning, replacement of spare parts)
major failures will require action by VXD crew

## Slide 21
Conclusions
Parts list almost complete, ordering/delivery has started
Construction to begin next month
Commissioning in fall 2015 (help by CERN colleagues)
Certification by German TÜV
Transport to KEK as a unit by summer 2016
Location of IBBelle on the backside of Tsukuba Hall
(climatized hut needed, will be provided by KEK)
Documentation in Twiki (started)