## Slide 1: Readiness of the FOS monitor for DESY’s thermo-mechanical mock-up. Kloster Seeon, May 19th DEPFET Workshop
D. Moya, I. Vila, A. L. Virto., J.González
Instituto de Física de Cantabria (CSIC-UC)
G. Carrión, M. Frövel.
Instituto Nacional de Técnica Aeronautica (INTA)

## Slide 2: Outline
Status and issues concerning the FOS integration in DESY’s thermo-mechanical mock-up.

## Slide 3: Fiber cables and sensors for the mock-up
Four cables with four sensors
Additional FOS with not specific support (see next table)

## Slide 4: Mock-up Fibers routing
For humidity fibers the routing will be the same as for the final solution
Teflon protected fibers( temperature) routed above the ring (The diameter is not small enough)

## Slide 5: FOS at DESY test beam
Fibres for PXD thermal set-up |  |  |  | 
Fibre Nº | Nº of sensors | position | protection | parameter
51 | 4 | On top of PXD on air | Teflon | Temperature
52 | 4 | On top of PXD on air | Teflon | Temperature
53 | 4 | On top of PXD on air | Teflon | Temperature
37 | 2 | On top of PXD on air | - | broken
38 | 4 | On top of PXD on air | - | Temperature + humidity
OS-4210 | 1 | To be defined | metal | Temperature
sensor 1 | 1 | To be define | Teflon | Temperature
Fibre 1 polyimida | 2 | To be define | - | Temperature + humidity
Fibre 2 polyimida | 2 | To be define | - | Temperature + humidity
SN89356/0003 | 1 | To be defined | metal | Temperature
 | New sensors |  |  | 
 | January 2015 Test Beam sensors |  |  | 

## Slide 6: Mock-up: Basic Fiber layout
Three fibers mounted inside a sealed Teflon tubing (0.9 mm ext. diameter) to prevent humidity sensitivity of the FBG sensors.
Additional fiber sensible to humidity partially inside a Teflon or steel tubing outside the sensitive volume.
Temperature sensitivity of the FBGs in the fibers from calibrations
The strain-induced temperature offsets to be determined just after the fiber fixation in the mock-up (requires precise temperature and humidity measurement at the fixation time using reference sensors).

## Slide 7: Protected fibers thermal calibration
Test beam fibers: same calibration parameters used
All the new fibers calibrated with temperature and the temperature sensitivity have been obtained.(s < 10 pm ( 1ºC) for almost all sensors)

## Slide 8: Strain reproducibility test
Check the repeteability of the FBG sensors readout during several fibre fixing test (3 times) .
Teflon protected fibers
Temperature reference sensor
Thermohygrometer
Fibers locking points

## Slide 9: Strain reproducibility test results
Very good repeteability ( s < 5 pm)
 |  | Mean ofset (nm) | s (pm)
Fibra 51 | F51-S1 | 1530,0027 | 3,4
 | F51-S2 | 1540,1130 | 1,4
 | F51-S3 | 1549,8907 | 3,2
 | F51-S4 | 1559,8850 | 4,8
Fibra 52 | F52-S1 | 1529,9430 | 3,9
 | F52-S2 | 1540,0732 | 2,3
 | F52-S3 | 1549,9356 | 2,7
 | F52-S4 | 1559,6151 | 1,3
Fibra 53 | F53-S1 | 1529,9510 | 4,8
 | F53-S2 | 1540,1448 | 1,0
 | F53-S3 | 1549,9159 | 2,0
 | F53-S4 | 1559,9192 | 2,1

## Slide 10: R/O of the fibers
For PXD mock-up DAQ ready: Readout integrated in epics similar to the test beam ( it is expected to be integrated smoothly)
The SVD fiber monitor will use the same interrogating unit. How should be proceed with the read integration ?

## Slide 11: Plans
Next week: final check on the kind of tubing to be used for the humidly fibers (Teflon or steel)
Assuming installation of fibers in the PXD mock-up during the first half of June