## Slide 1: SVD Sensors Update
Luciano Bosisio, Irina Rashevskaya
INFN - Trieste
8th Belle II VXD Workshop
Trieste, September 9-11, 2015

## Slide 2: Retesting of Scratched Micron Sensors
36 scratched sensors sent to Trieste end of June for retesting the P-side metal (AC) strips
  Standard AC strip scan on P-side
  Total sensor I-V characteristic for a quick check of possible (but not likely) damage to junctions
Retesting completed by mid-July
Metal shorts due to scratches separated in the first week of August
  91 out of 99 shorted strips separated
  of the remaining 8:
    4 in a sensor sent to Pisa in July
    4 non-separable (shorted by poly residue)

## Slide 3: Retesting of Scratched Micron Sensors
Priority | Sensor | Tested | Broken AC | Metal Shorts | Seen in Previous test | MetalShorts After Separation | Nr. of AC Defective Strips P-side | AC Defect %
1+ | 3077-15_Wedge | Vienna | 4 | 4 | 1 of 4 Broken AC | 4 | 8 | 1.0%
1+ | 3084-20_Wedge | Vienna | 0 | 0 | ? | 0 | 0 | 0.0%
1 | 3077-25_Wedge | Vienna | 6 | 6 | 3 of 6 Broken AC | 0 | 6 | 0.8%
1 | 3066-16_Wedge | Vienna | 5 | 0 | 4 of 5 Broken AC | 0 | 5 | 0.7%
1 | 3077-6_Wedge | Vienna | 5 | 2 | 1 of 5 Broken AC |  | 5 | 0.7%
1 | 3066-10_Wedge | Vienna | 0 | 2 | ? | 0 | 0 | 0.0%
1 | 3077-10_Wedge | Vienna | 10 | 0 | 2 of 10 Broken AC | 0 | 10 | 1.3%
1 | 3083-12_Wedge | Vienna | 2 | 0 | ? | 0 | 2 | 0.3%
1 | 3049-20_Wedge | Vienna | 10 | 3 | ? | 0 | 10 | 1.3%
1 | 3077-5_Wedge | Vienna | 5 | 3 | 5 of 5 Broken AC | 0 | 5 | 0.7%
1 | 3084-14_Wedge | Vienna | 1 | 2 | 1 of 1 Broken AC | 0 | 1 | 0.1%
1 | 3067-4_Wedge | Vienna | 1 | 2 | 1 of 1 Broken AC | 0 | 1 | 0.1%

## Slide 4: Retesting of Scratched Micron Sensors
2 | 3084-5_Wedge | Vienna | 2 | 2 |  | 2 | 4 | 0.5%
2 | 3066-8_Wedge | TS | 1 | 0 | No Added Defects | 0 | 1 | 0.1%
2 | 3084-12A_Wedge | TS | 4 | 0 | 1 new Broken AC | 0 | 4 | 0.5%
2 | 3067-10_Wedge | TS | 3 | 2 | No Added Defects | 2 | 5 | 0.7%
2 | 3083-1_Wedge | TS | 2 | 0 | No Added Defects | 0 | 2 | 0.3%
2 | 3084-7_Wedge | TS | 1 | 0 | No Added Defects | 0 | 1 | 0.1%
2 | 3084-6_Wedge | TS | 1 | 0 | No Added Defects | 0 | 1 | 0.1%
2 | 3067-7_Wedge | TS | 5 | 0 | No Added Defects | 0 | 5 | 0.7%
2 | 3078-10_Wedge | Vienna | 2 | 15 | 0 Broken AC , 0 ALShorts | 0 | 2 | 0.3%
2 | 3078-6_Wedge | Vienna | 1 | 8 | 0 Broken AC , 0 ALShorts | 0 | 1 | 0.1%
3 | 3084-8_Wedge | TS | 1 | 0 | No Added Defects | 0 | 1 | 0.1%
3 | 3078-13_Wedge | TS | 1 | 0 | No Added Defects | 0 | 1 | 0.1%
3 | 3066-20_Wedge | TS | 1 | 0 | No Added Defects | 0 | 1 | 0.1%
3 | 3083-10_Wedge | TS | 6 | 0 | No Added Defects | 0 | 6 | 0.8%
3 | 3084-17_Wedge | TS | 1 | 0 | No Added Defects | 0 | 1 | 0.1%
Priority | Sensor | Tested | Broken AC | Metal Shorts | Seen in Previous test | MetalShorts After Separation | Nr. of AC Defective Strips P-side | AC Defect %

## Slide 5: Retesting of Scratched Micron Sensors
Priority | Sensor | Tested | Broken AC | Metal Shorts | Seen in Previous test | MetalShorts After Separation | Nr. of AC Defective Strips P-side | AC Defect %
3 | 3084-19_Wedge | Vienna | 4 | 0 | 4 of 4 Broken AC | 0 | 4 | 0.5%
3 | 3083-15_Wedge | Vienna | 7 | 2 | 7 of 7 Broken AC | 0 | 7 | 0.9%
3 | 3067-20_Wedge | TS | 1 | 0 | No Added Defects | 0 | 1 | 0.1%
3 | 3048-3_Wedge | Vienna | 15 | 2 | ? | 0 | 15 | 2.0%
3 | 3049-3_Wedge | Vienna | 12 | 3 | ? | 0 | 12 | 1.6%
3 | 3067-16_Wedge | Vienna | 5 | 9 | ? | 0 | 5 | 0.7%
3 | 3067-25_Wedge | Vienna | 6 | 26 | ? | 0 | 6 | 0.8%
3 | 3077-13_Wedge | Vienna | 4 | 4 | ? | 2 | 6 | 0.8%
3 | 3077-16_Wedge | Vienna | 3 | 2 | ? | 0 | 3 | 0.4%

## Slide 6: Sensor measurements at HEPHY
No probecard
  instead three needles and stepping the table
Switching matrix to connect measurement devices
  Electromenter, SMUs, LCR-Meter
Automated measurement procedure
  LabView
(Manfred Valentan)

## Slide 7: Quality Test Center (QTC)
(Manfred Valentan)

## Slide 8
(Manfred Valentan)

## Slide 9: Analysis of HEPHY Sensor Test Data
Written in Python
  Input: raw data file from the QTC
  Output: csv file with list of defective strips, compatible with HephyDB
Combination of absolute and relative thresholds
  Absolute: e.g. current through dielectric -> pinhole
  Relative: e.g. doubled current (inside a certain deadband) -> implant short
Robust least-median-of-squares fit
  To find the expected value of a measurement outlier
(Manfred Valentan)

## Slide 10: Example: pinholes
(Manfred Valentan)

## Slide 11: Example: implant short
(Manfred Valentan)

## Slide 12: Analysis: Status and outlook
Progressing to final version
  The specifics of the measurement procedure are understood
  The thresholds are reasonable
  The analysis procedures are sophisticated
To do
  Understand a few special cases which are not caught by the present analysis procedure, but being pragmatic...
  Run the analysis over the data of all sensors
(Manfred Valentan)

## Slide 13: P-side contact problem effect on CAC measurement
P-side of Wedge sensors affected by poor metal-implant contacts
=> high effective series Rseries in Vienna measurement of CAC
=> the parallel R-C model underestimates C
admittance: Y = 1/Z
  Trieste measurement of CAC not affected by this problem, because the DC pad is not used.

## Slide 14: Metal to p-implant contact
Forward-bias I-V between p-strip and n-bulk, contacted by n-side Bias Ring
Likely some oxide present on the contact area: ‘burned’ by forcing 10 mA
The V-drop ∆V through the contact can be estimated from the curve shift
Chart: BL3066-19 P510 DCR Forward I-V
∆V

## Slide 15: Contact influence on AC-coupling cap. measurement
Even at 100 Hz and Vosc. = 0.5 V the parallel model underestimates CAC
The effect can be higher on other strips or sensors
Chart: f (Hz)
| 100.0 | 125.839 | 158.562 | 199.468 | 251.046 | 316.456 | 398.089 | 500.0 | 631.313 | 791.139 | 1000.0 | 1250.0 | 1578.95 | 2000.0 | 2500.0 | 3157.89 | 4000.0 | 5000.0 | 6315.79 | 7894.74 | 10000.0 | 12631.6 | 15789.5 | 20000.0 | 25000.0 | 31578.9 | 40000.0 | 50000.0 | 62500.0 | 80000.0 | 100000.0
Cp (Vosc=0.1 V) | 1.1425e-10 | 9.5308e-11 | 7.6989e-11 | 5.9558e-11 | 4.4476e-11 | 3.2235e-11 | 2.295e-11 | 1.6077e-11 | 1.1144e-11 | 7.8075e-12 | 5.4271e-12 | 3.9317e-12 | 2.897e-12 | 2.2207e-12 | 1.7947e-12 | 1.5044e-12 | 1.3014e-12 | 1.1813e-12 | 1.0687e-12 | 1.0096e-12 | 9.4481e-13 | 9.1569e-13 | 8.9117e-13 | 8.7213e-13 | 8.5937e-13 | 8.4653e-13 | 8.3284e-13 | 8.1819e-13 | 8.0064e-13 | 7.7659e-13 | 7.4974e-13
Cp (Vosc=0.3 V) | 1.4574e-10 | 1.3556e-10 | 1.2492e-10 | 1.1325e-10 | 1.007e-10 | 8.7345e-11 | 7.483e-11 | 6.1662e-11 | 4.8923e-11 | 3.8726e-11 | 2.8819e-11 | 2.1361e-11 | 1.5683e-11 | 1.1283e-11 | 8.5923e-12 | 6.5263e-12 | 5.1106e-12 | 4.2379e-12 | 3.3203e-12 | 2.8515e-12 | 2.1484e-12 | 1.78e-12 | 1.5819e-12 | 1.4321e-12 | 1.3114e-12 | 1.1592e-12 | 1.0527e-12 | 9.5938e-13 | 8.8322e-13 | 8.2671e-13 | 7.7768e-13
Cp (Vosc=0.5 V) | 1.617e-10 | 1.5522e-10 | 1.4821e-10 | 1.4015e-10 | 1.3105e-10 | 1.2055e-10 | 1.0934e-10 | 9.6671e-11 | 8.2689e-11 | 6.9451e-11 | 5.6835e-11 | 4.4919e-11 | 3.455e-11 | 2.5991e-11 | 2.0166e-11 | 1.5443e-11 | 1.2291e-11 | 1.0185e-11 | 8.166e-12 | 7.0495e-12 | 5.819e-12 | 4.6619e-12 | 3.7894e-12 | 3.2828e-12 | 2.8428e-12 | 2.3937e-12 | 2.1412e-12 | 1.9099e-12 | 1.6394e-12 | 1.4246e-12 | 1.2477e-12
Cs (Vosc=0.1 V) | 1.8942e-10 | 1.8801e-10 | 1.8652e-10 | 1.8435e-10 | 1.8236e-10 | 1.7974e-10 | 1.7644e-10 | 1.719e-10 | 1.651e-10 | 1.5603e-10 | 1.436e-10 | 1.2967e-10 | 1.1338e-10 | 9.4323e-11 | 7.6447e-11 | 5.8156e-11 | 4.3037e-11 | 3.1339e-11 | 2.1989e-11 | 1.5456e-11 | 1.0673e-11 | 7.4058e-12 | 5.2831e-12 | 3.7123e-12 | 2.7535e-12 | 2.0882e-12 | 1.6781e-12 | 1.4259e-12 | 1.2472e-12 | 1.1157e-12 | 1.0238e-12
Cs (Vosc=0.3 V) | 1.8867e-10 | 1.8748e-10 | 1.8628e-10 | 1.8461e-10 | 1.8323e-10 | 1.8136e-10 | 1.7969e-10 | 1.7824e-10 | 1.7725e-10 | 1.7634e-10 | 1.7228e-10 | 1.6945e-10 | 1.6433e-10 | 1.5543e-10 | 1.439e-10 | 1.2734e-10 | 1.0929e-10 | 9.2012e-11 | 7.9129e-11 | 6.5497e-11 | 5.6831e-11 | 4.6534e-11 | 3.7871e-11 | 2.8397e-11 | 2.1589e-11 | 1.5742e-11 | 1.1835e-11 | 8.6657e-12 | 6.2441e-12 | 4.6598e-12 | 3.5758e-12
Cs (Vosc=0.5 V) | 1.8912e-10 | 1.8827e-10 | 1.8745e-10 | 1.8651e-10 | 1.8539e-10 | 1.8379e-10 | 1.8199e-10 | 1.8006e-10 | 1.7878e-10 | 1.7812e-10 | 1.7405e-10 | 1.7106e-10 | 1.6681e-10 | 1.6032e-10 | 1.5051e-10 | 1.3735e-10 | 1.2205e-10 | 1.0676e-10 | 9.2675e-11 | 8.0109e-11 | 6.7351e-11 | 5.9362e-11 | 5.5452e-11 | 4.8039e-11 | 4.0786e-11 | 3.4093e-11 | 2.7957e-11 | 2.3279e-11 | 1.9011e-11 | 1.54e-11 | 1.259e-11
↑
Vienna test f (100 Hz, Vosc. = 0.5 V)
Measurement of CAC versus f at different test voltages Vosc.
(performed in Trieste on a sample strip)

## Slide 16: Contact influence on AC-coupling cap. measurement
This is a rather extreme case, with very high contact resistance on the last strips
Triested measurement
Chart: Strip number
Chart: Strip number
Micron Sensor 3084-20 – Vienna test data
CP
CS

## Slide 17: Contact influence on AC-coupling cap. measurement
On most sensors the series model improves the quality of the CAC measurement
Micron Sensor 3066-10 – Vienna test data
Chart: Strip number

## Slide 18: Test of a Large HPK Sensor
Front-side Rings are bonded to these traces
Sensor # 83 (not classified) recently tested in Trieste

## Slide 19: Sensor testing set-up

## Slide 20: Sensor testing set-up
Front and Back-side Rings are connected through cables
Strips are contacted by a 40-needle Probe Card

## Slide 21: Test results on Large HPK Sensor
P-side:
  1 AC defect (pinhole)
  2 (adjacent) high-current strips (0.53 µA, 0.71 µA)
  4 implant opens (verified at the microscope)
  1 implant short to adjacent non-readout strip
N-side:
  No AC defects
  1 high-current strip (0.75 µA)
  24 implant opens (verified at the microscope)
=> Consider as Class A ?
We can test some more if available