## Slide 1: DHH at DESY Test Beam 2016
Igor Konorov
TUM Physics Department E18
19-th DEPFET workshop
May 11-15 2016
Kloster Seeon
Overview:
  DHH system overview
  DHE/DHC functionality
  Problems and fixes
  Plans

## Slide 2: Simplified DHH system at DESY
DHC
DHE
DHE
ONSEN
DAQ
FTSW
EPICS
Trigger
127MHz
Bonn
DAQ
Complex system
Requires multiple tests for verification and debugging
We provide additional features for tests

## Slide 3: ONSEN standalone tests
DHC
DHE
DHE
ONSEN
DAQ
FTSW
EPICS
Trigger
127MHz
Bonn
DAQ
For every trigger DHC sends UDP packet with HLT information for ONSEN, standard firmware feature
DHC generates predefined hit pattern, special firmware
PC

## Slide 4: DHH control and monitoring
DHC
DHE
DHE
ONSEN
DAQ
FTSW
EPICS
Trigger
127MHz
Bonn
DAQ
Bonn DAQ - very important tool
  Pedestal determination in local run mode
  DQM during data taking
  Generate log file of data consistency check

## Slide 5: DHE data flow
No DDR memory
Simple event definition
  DHP data frame defines start of event,
  Timeout defines end of event
DHP 0
FIFO
Event
Builder
DHP 1
FIFO
DHP2
FIFO
DHP 3
FIFO
Aurora
DHC
DHE
DHE data flow diagram
MX

## Slide 6: DHC data flow
DHE 0
FIFO
256 MB
Event
Builder
DHE 1
FIFO
256 MB
DHE2
FIFO
256 MB
DHE 3
FIFO
256 MB
Aurora
ONSEN
DHC
DHC data flow diagram
MX
DHE 4
FIFO
256 MB
FIFO
256 MB
UDP
Bonn DAQ
DHC event builder actions:
  Trigger initiates generation of DHC Start Of Event frame
  Reads DHE0 FIFO until DHE End Of Event frame
  Reads DHE1 FIFO until DHE End Of Event frame
  Generates DHC End Of Event frame
Trigger #

## Slide 7: Readout stability
Setup was running at 65 MHz and 10m Infiniband cable
  IB => DHE0 – DHP links unstable
  OB => DHE1 – all 4 DHP links stable
  Observed event number mismatch between DHE0 and DHE1 & DHC event numbers. DHE0 event number is always smaller
DHE0 was exchanged by another one
  DHE0 - DHP links became more stably
  System was running without loosing synchronization
There was no time to measure Bit Error Rate
Preliminary conclusion:
Links instability caused “extra event” in DHE and loss of synchronization at DHC event builder
Why links became more stable after exchanging DHE?

## Slide 8: Summary of Observed DHH problems
 
  No memory dump support for pedestal determination
  Limited DHE FPGA memory buffer, DDR buffer was not included in DHE yet
 
  Problem of handling big events
  No FPGA logic to handle unforeseen data rate/event sizes
  Limited buffer sizes
 
  Event number mismatch at DHC event builder
  Unstable DHP links cause data errors and/or loss of link
  No recovery logic neither in DHE nor in DHC to handle wrong Data Format
  Other reasons ? To be investigated

  65MHz DHP clock instead of 76 MHz => Asynchronous operation

## Slide 9: Observed DHH problems
 
  Bonn DAQ error logs: timeout error between frames within one event(>5ms)
  DHH features multiple data buffers including DDR memory. DDR interface is optimized for bandwidth on expense of bigger latency

## Slide 10: Improvement of DHH functionality
DHE: DDR memory for data buffering
  Sharing detector data between overlapping events
DHE: DHP data consistency check and recovery logic
DHC: DHE data consistency check and automatic resynchronization
DHE: include programmable pixel remapping
DHE: include clustering algorithms
DHC: internal trigger generator for local runs

## Slide 11: Conclusions
System was running stably with some problems
PXD is complex system
Such system requires availability of all components and significant time for debugging. Time was limited.
New diagnostic tools and information being included during system test but more is needed (ex.:event number mismatch)
Data consistency check and automatic resynchronization to be included in all steps of data processing
The DHH worked stably an no principle problems were observed

## Slide 12: Instead of summary
DHE and DHC firmware still require a lot of development
We will try to work out a schedule to complete all developments by end of September

## Slide 13: Hardware development plans
ATCA carrier card
  If display port cable fulfils requirements present the final system layout to TB in June
  Evaluation of new optical interfaces including irradiation test - June - July
  Production of final ATCA carrier card for end of Septemebr

## Slide 14
THANK YOU