## Slide 1: Gemba Summary
Plan: Shipment (sea transport in summer) 6-8 weeks => arrival August/September
Commissioning: After October B2GM
Needed for Beast Phase II: April 2017
Visit at KEK by H.-G. Moser, R. Sedlmayer, and S. Vogt to prepare IBBelle commissioning (April 23-26)
Check location, service connections, logistics

## Slide 2: Location
UNIT A
UNIT B (laterl)
CO2 line
Pump removal
Enough service space?
Room next to Tsukuba Hall in B1 level (used for solenoid)

## Slide 3: Requirements
60 cm free space around frame for access (open service doors).
Sufficient space for replacement of pumps (lift and crane access).
Short CO2 transfer lines.
Possibility to fit second unit later without displacing the first one (clearance for crane operation).
Don’t obstruct escape doors.
60 cm
60 cm
Remove pump
2 m
Need fork for crane
1 m

## Slide 4: Location at KEK
2.1 m
0.6 m
1.5 m
>0.6 m
For solenoid
Clearance?
Unit B
Unit A

## Slide 5: Installation of B-Unit
0.6 m
1.5 m m
>0.6 m
For solenoid

## Slide 6: Escape Door
Escape door hidden behind unit B
60 cm clearance between wall and IBBelle frame
(+ 40 cm on frame)
In principle we can short Unit B by ~140 cm
(no accumulator needed)
Only partially hidden.
Is this acceptable ?
Seems to be ok….

## Slide 7: Ramp
In order to extract IBBelle the container has to be levelled and some steps filled
(slight ramp before gate)
Request to transport company(?)

## Slide 8: Long Transfer Line
At IBBelle Location
Along gallery (moveable)
Total length: 32 m (as expected)

## Slide 9: Transformer
Voltage too high (420V instead of 400V)
Might reduce lifetime of some components
Phase shifts between the lines were not correct!
Fixed?
Still confusing measurements
=> Need to verified

## Slide 10: Other Services
Compressed air (6 bar, ok)
flexible tube (standard) will be provided by MPP
sample of connection pipe at KEK
Water pipe connection (25°C, 3 bar)
Flexible hose sufficient. Sample of connection type at KEK.
Network
Lan required, power for switch close to JB (Ito-san)
CO2 connection IBBelle  transfer line  JB
  VCR, 3/8 and 1/2 inch. Needs to be produced in Japan
CO2 Connection JB <-> manifolds:
  VCR, 3/8 and 1/2 inch. Needs to be produced in Japan. Need tray for tubes and cables (to give headroom).
Vacuum for Flex lines:
10-7 bar => ok
IBBelle shipment in summer. After locating it in B1 the exact pipe shapes and length can be defined. Time for production and installation: till October 2016

## Slide 11: Other Requests
Monitoring of environment (humidity, temperature):
  Dew point < 15°C.
R404a filling
By company. Quantity: 12 kg. Needed early October
CO2 filling:
by us or KEK group, need 3 bottles 50kg (with rising pipe) in B1
Unloading IBBelle from container:
  need to level container

## Slide 12: Tools needed
Crane tool: bar
Fork to lift pumps with the crane (only if unit B installed)
5 m x 2 m; 5 tons

## Slide 13: Additional budget for IBBelle
Shuji Tanaka:
AC transformer (400V: 3phase with neutral line)
Water Cooling pipe
Vacuum pump for flex line
Transfer tube piping (also between junction box and manifold)
CO2 monitor
Dew point sensor for IBBelle room