## Slide 1: Gated Mode Timing Observations

## Slide 2: The Gated Mode
SwitcherB allows to operate Gating in two different Modes:
Gated Mode without Readout: Clocking is disabled, Clear moves to high for all gates except current one
  Needs fast clocking to synchronize with correct readout position
  Problem with DHPT1.0: when gating starts at end of frame, SerIn signal is omitted because of missing clock and hence full frame will be lost. Hopefully solved with new DHPT1.2 version.
Gated Mode with Readout: Clock continues to run, so frame readout will be synchronized with SerIn, Clear changes to high level immediately on non-active channels
  Losing at least 600ns for turning in & out of Gated Mode
  Hope for lower pedestal oscillations

## Slide 3: Gated Mode Simulation 11 gates @250MHz = 1,408µs
At least 5 StrG for initiating GM
36ns for Switching all Switcher Groups ON or OFF
Pure GM 1124ns
Fast
Clocking &
Returning to
Normal
Operation
224ns
1 gate @250MHz = 128ns
Disable GM:
switch StrC to low +
falling edge of StrG

## Slide 4: Oscilloscope Sw1/1 vs Sw6/4 GM 11 gates
Sw6/4
800ns ramp-up time
2.3µs total time
>1µs longer than
first Switcher group
Rise time Sw1/1 ~ 200ns
Rise time Sw6/4 ~ 800ns
This compares to a substantial lower rise time in the SwitcherB manual but bear in mind that 6 SwitcherBs, separated into 4 groups@8Clear channels have to be switched on at once (4 times consecutively within 36 ns)
Only first and last Clear contact are accessible via oscilloscope

## Slide 5: Laser Scans GM 11 gates sequence
Full Protection
~600ns
Laser Activation Timing moved forward in 10ns steps
~200ns

## Slide 6: Pedestal Oscillations: 11 gates GM-sequence
GM starts at gate 24*
* DHPT1.0 Bug: GM is programmed for gate 22 but normal mode repeated twice
GM Gate 25-33 Clear High
Gate 34:
at least 5 StrG
signals finish GM
Gate 35-36:
Fast Clocking
Higher pedestals due to Sw6/4?
Extreme Amplitudes delayed reaction of Switchers 200-800ns
~2-6 gates
Total GM period of 18 gates including pedestal variations

## Slide 7: Oscilloscope Sw1/1 vs Sw6/4 GM 11 gates zoomed
Time base: 1µs/div
gate 24
gate 28
gate 34
gate 42

## Slide 8: Gated Mode Simulation 3 gates @250MHz=384ns
At least 5 StrG for getting into GM
36ns for Switching all Switcher Groups ON or OFF
Pure GM 228ns
Fast
Clocking &
Returning to
Normal
Operation
84ns

## Slide 9: GM11gates vs GM3gates first Clear
Measured 14V
(Chigh=20V, Clow =5V)
GM protection
effective @ 9V
above Clow

## Slide 10: GM11gates vs GM3gates last Clear

## Slide 11: GM short 384ns sequence
pure GM period reduced by 7 gates,
fast clocking now only 1 gate
everything else equal
Total GM period reduced to 10 gates = 1,28µs including pedestal variations
STRANGE:
Reason for higher pedestals 4 gates after stopping GM unclear
same starting sequence pedestal peak at gate 26

## Slide 12: It’s not only a time constant… 250 MHz
Total GM period of 10 gates including pedestal variations

## Slide 13: Pedestals back in range at gate 34 @285 MHz
Total GM period of 10 gates including pedestal variations

## Slide 14: Short Cables – 11 Gates GM w/o RO overview
No idea where this is coming from
DHPT reset at
end of frame

## Slide 15: Short Cables – 11 Gates GM w/o RO zoomed

## Slide 16: Long Cables – 11gates GM w/o RO overview

## Slide 17: Long Cables – 11gates GM wo RO zoomed

## Slide 18: 3gates GM w/o RO overview (DCD3 short cables)

## Slide 19: 3gates GM w/o RO zoomed (short cables)

## Slide 20: 6gates GM with Read-Out overview

## Slide 21: 6gates GM with RO zoomed
although relatively short GM period pedestal oscillations persist longer than even 11gates w/o RO
variation spread thereafter seems to be substantially lower!

## Slide 22: 10gates GM with RO overview

## Slide 23: 10gates GM with RO zoomed
Pedestal oscillations are within arbitrarily chosen threshold limit of ±5 ADUs
not before 23 gate periods = 2,94µs

## Slide 24: GM11gates w/o RO analog CMC overview

## Slide 25: GM11gates w/o RO analog CMC zoomed
In general ACMC looks better but risk that leveling out absorbs signal charge information

## Slide 26: ACMC, Single Row Pedestal Oscillations
Distinct DCD-Gradient disappeared after fixing cooling problem and attaching additional capacities on dock-box and break-out board
Low IPaddIn seems to “swallow” low signal charge
BEWARE: ACMC seems to amplify higher signals

## Slide 27: Analog CMC, IPAddIn = 30

## Slide 28: Analog CMC, IPAddIn = 60

## Slide 29: Summary (1)
Fast switching in & out of Gated Mode is key for the performance of the pixel detector
Various parasitic capacitances extend
rise and fall times of Clear voltage
at gate level
For various reasons (reducing deadtime, sync with CDC) the BELLE II collaboration wants to keep the total GM period (GM + pedestal oscillations) as short as 1µs
Pedestal Oscillations after the Gated Mode can prevent the detector from recording reasonable data
The presented analysis is based on meager Switcher Substrate connection at least for Switcher 5 + 6 (PXD9 Pilot Module), hence it is very difficult to draw accurate conclusions.

## Slide 30: Summary (2)
All we have are the following observations:
  Pedestal Oscillations constitute a complex process, not everything can be traced back to a time constant
  Pedestal Spikes after finishing GM look strange and have potentially something to do with the long fall time for the last switchers
  It was shown that the shortest GM sequence (3gates w/o RO) reduces total GM time to 10 gates which would be equivalent to ~1µs @305MHz – CAUTION because of last switcher
  GM with readout seems to elongate the total GM period although once falling under a certain threshold level pedestal variations appear much smoother
  Mounting additional capacities on the dock-box and the patch panel as well as the analog CMC seem to reduce the Pedestal Oscillations
Looking forward to PXD9 series 7 incorporating heavily improved switcher voltage supply lines in order to find out more resilient results

## Slide 31
Thank you for your attention!

## Slide 32: Analog CMC, IPAddIn = 60, ECG