## Slide 1
DCD waveform measurements
-
DHE DAC & DCD ADC gain
-
Offset strength measurements
F. Lütticke, B. Paschen
On behalf of the Lab and Test Beam Groups

## Slide 2: DCD communication
VDD
VDD
VDD
GND
GND
DOx[y]
floating
Transmitter
VDD
VDD
VDD
GND
GND
VREF
Reference
D
!D
!D
D
DCD
DHP
x64
x1
x64
Receiver
VDD
VDD
GND
GND
Can be overwritten from the
Outside (this measurement!)
DCDBpp 3: Uses VDDD
For DCDB 4.1/2: Planed to connect reference to to VDDA. Done?

## Slide 3
Low VREF: Always receive logical 1
High VREF: Always receive logical 0
Results (DCDBpp + DHPT1.0)

## Slide 4
Extract voltage threshold from data to get absolute voltages

## Slide 5
All bits of one link at 1.9V VDDD DCDpp 3, 250 MHz

## Slide 6
Low VREF: Always receive logical 1
High VREF: Always receive logical 0
Results (DCD4.1 + DHPT1.1)
OLD
Less noise on line
“faster” high/low transition

## Slide 7
DCDB4.1 lines stacked

## Slide 8
All bits, all links at 1.9V VDDD DCD 4.1, boost on, 305MHz

## Slide 9
Comparison
OLD, 250 MHz
~300mV
~400mV

## Slide 10
Comparison (cont’d)
~150mV
~300mV
New, 305MHz
1 link
New, 305MHz
all links
OLD, 250 MHz

## Slide 11
DCDB4.2 lines stacked
LSB
MSB

## Slide 12
All bits, all links at 1.8V VDDD DCD 4.2, boost on, 305MHz
Artifacts due to low noise and large steps

## Slide 13
1.8V VDDD DCD 4.2, 305MHz, Boost comparison
Boost On
Boost Off

## Slide 14
Offset of around 20uA
resisitivity 226.8 kOhm
-3.827 +- 0.006 nA/DAC
DHE current source – new DHE

## Slide 15
Offset of around 0.05uA
resisitivity 135301.1 kOhm
-3.814 +- 0.000 nA/DAC
DHE current source – old DHE

## Slide 16
gain of H5.0.26 – DCDB4.2
Rf [kOhm] | gain [nA/ADC] | range [uA] | range (+-100 ADC) [uA]
26 | 75 | 19 | 15
19 | 96 | 25 | 19
13 | 144 | 36 | 29
8.7 | 171 | 44 | 34
7.7 | 217 | 61 | 48
6.0 | 313 | 80 | 63

## Slide 17
gain of H5.0.26 – DCDB4.2
 | DCDB4.2 Bonn | DCDB4.1 Göttingen | DCDB4.2 Bonn | DCDB4.1 Göttingen
Rf [kOhm] | gain [nA/ADC] | gain [nA/ADC] | range [uA] | range[uA]
26 | 75 | 67 | 19 | 17
19 | 96 | 88 | 25 | 22
13 | 144 | 128 | 36 | 33
11 | 171 | 156 | 44 | 40
8.7 | 217 | 192 | 56 | 49
7.7 | 239 | 213 | 61 | 55
6.0 | 313 | 278 | 80 | 72
luetticke@uni-bonn.de

## Slide 18
Offset Calibration (DCD4.2 + DHPT1.1)

## Slide 19
Offset Optimization(DCD4.2 + DHPT1.1)

## Slide 20
Offset strength (DCD4.2 + DHPT1.1)
0.05LSB/DAC
2.35LSB/DAC
5.87LSB/DAC
8.57LSB/DAC
1.2uA/DAC 3.6 x I0
0.84uA/DAC 2.4 x I0
0.34uA/DAC 1 x I0
0 uA/DAC
Strength
Range
152 uA
106 uA
43 uA
Design value
180 uA
120 uA
60 uA