## Slide 1: Wide Range Charged Particle Detection Using Fast Scintillator and Si-PM
El-Mul Technologies, Ltd.
Amit Weingarten, Ph,D, R&D Manager
18-Oct-2017
Notes: 1

## Slide 2: Outline
Who are we?
Challenges of electron and ion detection
Sensor comparison
Si-PM as sensor in SEM and TOF systems
Si-PM feasibility study: dynamic range, stability, timing simulations
Summary and future needs

## Slide 3: El-Mul Technologies Ltd.
El-Mul is 25 years old, ~30 employees, privately owned company located in Rehovot, Israel
Specializes in design and manufacturing of detectors for charged particle detection in SEM, FIB (ion-electron beams), lithography, and Mass Spectrometry systems
Manufacturing >1000 detectors / year

## Slide 4: Challenges in Charged Particle Detection
Limited Space: BSE in SEM, Mass Spec, TOF SIMS
Fast detection: TOF, charging in SEM
Wide dynamic range: most systems require ≥3 orders of magnitude
Heat evacuation: in-vacuum systems
Lifetime: cost of ownerships

## Slide 5: Basic Requirements
Detection efficiency 100% (all particles reaching the detector will produce signal)
Two working regimes
  Pulse Counting mode: Iinput< few pA
  Analog (current) mode: 1 pA < Iinput<1 nA
Linearity over analog regime (typically ~1%)
Overall timing FWHM for pulse separation <5ns (Scintillator + sensor)
Envelope can be as small as few mm (at least in one dimension)
Power < 1 Watt

## Slide 6: Possible Sensors
 | Pros | Cons
PMT | Large dynamic range Linear | Large - complicated mechanics, complex LG/ fiber  low LCE Limited output current Heat generation
Si diode | High QE Thin Large dynamic range Linear | High capacitance – noisy and slow for large areas Low gain –additional amplifier
Si-PM | Small- simpler mechanics and LG  high LCE Lower overall cost Lower heat | Complex electronics Linearity (?) Gain & thermal stability (?)
MCP / Electron Multiplier | Fast High gain Linear | Low QE Short lifetime High voltage required

## Slide 7: Fast Scintillator
An enabler for using a fast light sensor is a fast scintillator
El-Mul developed ScintiFast: inorganic high-yield fast scintillator
  Photon yield: 50-200 photons / 10keV electron
  Decay time: <0.5ns to 1.5ns
  Trade-off between yield and decay time

## Slide 8: Backscattered Electron Detector
To achieve high BSE collection efficiency (CE) in SEM systems, detector is best placed between the column and the sample
Typically large area Si-diodes are used for high CE, but this results in high capacitance that limits the speed and produce noise
Suggested alternative: thin scintillator and light guides, coupled to few Si-PM.
  Advantage: high speed, high gain, low noise

## Slide 9: Backscattered Electron Detector
To achieve high efficiency and good uniformity, various material, geometrical designs and coatings were simulated
Light Guide
scintillator
Al coating
Si-PM

## Slide 10: Ion Detectors
Ion detectors are needed in FIB, Mass spec, TOF-SIMS
Single ion detection is required
When fast detection is needed, detectors are based on either MCPs, electron multipliers, or plastic scintillators (with PMT) which have limited lifetime
Other limitations are space and heat

## Slide 11: Example: Ion Detector
Required wide dynamic range from low rate (single ions) in pulse counting mode to high current up to ~1nA
Single ions up to ~3pA  Rate < 2x107 ion/sec ~50ns between pulses  need <5ns resolution to ensure no overlap
Analog mode 1pA-1nA with 1% linearity
Very limited space: Si-PM is natural choice due to small size and low heat consumption
Ions
Ion-Electron Convertor
Scintillator
(at HV)
Light Guide
Si-PM
Electronics

## Slide 12: Analog Regime
1 pA <Iion< 1 nA
Assuming: electron CE from convertor of 100%, scintillator yield of 50ph/e @5kV, Light CE of 33%, QE / PDE of 25% @ sensor, gain ~ 2.3x105
   1mA < IOUT< 1mA
Well above dark current level
For using PMT, one needs to work at much lower gain. Since amplifier gain is limited by bandwidth, additional amplification stage is needed
   cost and noise

## Slide 13: Si-PM Selection (3x3mm)
 | 10mm | 15mm | 25mm | Comments
Gain @ recommended Vb | 1.4x105 | 2.3x105 | 5.2x105 | 
PDE | 0.1 | 0.25 | 0.35 | 
# of pixels | 90,000 | 40,000 | 14,400 | 
Ph-e/e | 1.6 | 4 | 5.6 | For 50ph/e from scintillator; 33% LCE
Linearity (at max Iin) | 0.18% | 1% | 3.8% | 
Values are based on Hamamatsu datasheets
Linearity was estimated using:
Nfired=Ntotal [1-exp(-Nphoton*PDE/Ntotal)]

## Slide 14: Electron Collection From Convertor
Scintillator and convertor are few mm only
Electrons emitted from the middle and 4 corners of the converter
4-5kV are required so all electrons will be collected
Converter
Scintillator
incoming ions
12eV SE trajectories – top view
12eV SE trajectories – side view

## Slide 15: Si-PM Feasibility – Questions Addressed
Timing and electronics
Dark count – what is the threshold required to eliminate false detection?
Linearity
Gain and thermal stability

## Slide 16: Timing
With dedicated electronics that gives fast output pulse: FWHM ~1.5ns
Taken from SensL datasheet

## Slide 17: Electronics Design
Pre-Amplifier circuit was designed with small resistor as load to enable large bandwidth
Simulations were performed with 0.5ns input pulse that simulate input from scintillator. Pixel Capacitance 0.05pF
Predicted pulse decay is ~2ns
2ns per major tick

## Slide 18: Dark Events Distribution
With 15mm pixel – pulses were too low to characterize . 50mm pixels were used
According to Hamamatsu datasheet: 3% are double Ph-e events due to cross talk
Some triple Ph-e events are seen but 4 Ph-e events are very rare (many seconds)
Raising trigger level > 4ph-e – no triggering at all
4 ph-e threshold can be used giving negligible “false detection”
1 Ph-e
2 Ph-e
3 Ph-e
Afterpulses

## Slide 19: Stability Measurements Set-Up
LED source operated by function generator: variable frequency, amplitude, duty cycle
  LED stability is monitored in real time by Si-diode
Output reading by scope
Si-PM was placed on thermal pad placed on Cu plate

## Slide 20: Experimental Setup
Pulser-2
(T1, T2)
Pulser-1 (T3, T4)
Si-PM mount on Cu plate
Scope
LED
Lens
T1
T2
T3
T4
“Line” or “scan”
“Frame”
Idle time

## Slide 21: Bias Circuit Optimization
Bias circuit scheme is shown
We use 100W to minimize voltage drop of the bias voltage with high current (1mA=100mV)
Increasing the capacitance showed improved stability within “line”
  Charge in small capacitor too small to supply the required current during the “line” stage (up to 1mC)
At 100uF the signal becomes very stable
100W
100mF

## Slide 22: Linearity
Linearity was tested over a range of ~10 up to upper limit of analog regime
Deviations up to 6% but no trend – most likely result from inaccuracies

## Slide 23: Frame Stability – Measured Traces
At 20ms sampling rate, acquisition of 1.3s  43 full frames

## Slide 24: Stability within Frame
LED relative stdev of line-average within frame: 0.5%
Si-PM Relative stdev of line-average within frame: 1.2%

## Slide 25: Measurement Results Examples
Within acquisition: range of frame average variations ±1%
“Jump” at middle of the acquisition is from scope because it is seen also in LED signal in the same direction (positive) although the photodiode signal is negative
LED / Photodiode stability within acquisition is ~0.2% (before step and after step)
Frames 1-22 and 23-43 should be analyzed separately

## Slide 26: Variations Between Acquisitions
Stdev of frames average before “jump” and after “jump” is 0.5-0.7%
LED Stdev of frames is much smaller: 0.1-0.2%

## Slide 27: Stability Conclusions
Frame-to-frame variations over 1sec between 0.1-0.2% and 0.5-0.6% for different pulse durations, although in all cases LED show <0.2% variations.
Line-to-line variations ~1.2% (LED 0.5%)
Reason for difference is not clear. Measurement not taken consecutively – possibly stray light or other source of noise.

## Slide 28: Open Questions
Effect of cross talk at low side of analog regime
Effect of Gain temperature variation on SNR and stability
  Response to ambient temperature variations ~1.5% per degree
On bench no issue was observed, but in vacuum…?

## Slide 29: Summary
We demonstrated that for some SEM and Mass Spec configurations, Si-PM can be the sensor of choice offering the best solution for the system from both design and performance considerations
Si-PM was characterized to verify compatibility to system requirements:
  Dark count: >3 ph-e threshold can be used for pulse counting regime giving negligible “false detection”
  Linearity: linear over ~10 in the upper end of analog regime
  Stability over 1 hour, 0.1-1.2% variations
Pre-amplifier simulations support feasibility

## Slide 30: Future Needs
Larger Si-PM with small pixels – wider dynamic range
Lower dark current and cross talk
Higher PDE for lower pixels
  Also enables reducing scintillator high voltage
Faster, smaller pixels (but with reasonable PDE)

## Slide 31: Thank You!
Questions and Discussion
Notes: 31