## Slide 1
LIGHT 17
G. Bonanno – INAF O.A. Catania – LIGTH 17 – Ringberg Castle, 19 – October 2017
Characterization of recently manufactured LVR2 and LVR3 HPK SiPMs suitable for SST-2M telescope
Giovanni Bonanno – INAF Catania
for the ASTRI Collaboration & the CTA Consortium
Universidade de São Paulo
Instituto de Astronomia, Geofisica
e Ciencias Atmosferica

## Slide 2
Summary
G. Bonanno – INAF O.A. Catania – LIGTH 17 – Ringberg Castle, 19 – October 2017
Why 57.4mm x 57.4mm with 7x7 mm2 pixels
Characterized detectors
Characterization Setup
Optical Cross Talk evaluation of LVR2 7075
PDE at 405 nm vs OV measurements
PDE in the 285nm – 850nm spectral range
LVR3 vs LVR2 comparison
Effect of the IR filter on the OCT

## Slide 3
Characterized devices
G. Bonanno – INAF O.A. Catania – LIGTH 17 – Ringberg Castle, 19 – October 2017

## Slide 4
Why 57.4mm x 57.4mm with 7x7 mm2 pixels
G. Bonanno – INAF O.A. Catania – LIGTH 17 – Ringberg Castle, 19 – October 2017
8 x 8 pixels
7x7 mm2
FF= 93.8%
57.4 mm
57.4 mm
Minimum distance
for safety conditions
PDM coverage by assembling 7x7 mm2 chips with 0.2 mm interspace
FILLING FACTOR PER TILE:
93.8% !!!
Sensitive area:
6.975 x 6.975 mm2 (75 µm cell)
93 x 93 = 8649 micropixels

## Slide 5
Characterization Setup based on pulsed sources
G. Bonanno – INAF O.A. Catania – LIGTH 17 – Ringberg Castle, 19 – October 2017
Pulse Generator
LASER and LED Drivers
Reference Photodiode NIST calibrated
Integrating
Sphere
Citiroc
FEE
Calibrated Neutral
Density Filter

## Slide 6
Pulsed sources available at our lab
G. Bonanno – INAF O.A. Catania – LIGTH 17 – Ringberg Castle, 19 – October 2017
19 pulsed light sources in total are available at the COLD lab: 3 LASERs and 16 LEDs
They allow to cover the 285nm – 850nm spectral range

## Slide 7
OCT evaluation through Dark stairs
G. Bonanno – INAF O.A. Catania – LIGTH 17 – Ringberg Castle, 19 – October 2017
Electronics inside the ASIC for staircase measurements
Shaping time 15 ns
Fast Shaper
PZC inadequate
OCT versus temperature. Below 5°C, the OCT depends only by the SiPM size while above 5°C the OCT is affected by the pile-up effect. PZC inadequate..

## Slide 8
LVR2 7075 - OCT vs OV – CS and CN type
G. Bonano – INAF O.A. Catania – LIGTH 17 – Ringberg Castle, 19 – October 2017
The OCT values to be considered (close to the effective) are those at 2°C

## Slide 9
PDE measurements with pulsed sources
G. Bonanno – INAF O.A. Catania – LIGTH 17 – Ringberg Castle, 19 – October 2017
PHD for the MPPC 3x3 mm2 (micro-cell 75µm) flashed (left plot) and in dark (right)

## Slide 10
PDE @405nm vs OV: Comparison 7075 vs 7050
G. Bonanno – INAF O.A. Catania – LIGTH 17 – Ringberg Castle, 19 – October 2017
The PDE at 405nm of the 7075 is 14% greater than that of the7050 at all Over Voltages in the 2V - 6 V range.

## Slide 11
PDE in the 300nm – 850nm spectral range
G. Bonanno – INAF O.A. Catania – LIGTH 17 – Ringberg Castle, 19 – October 2017
At 3V of OV the PDE at the peak of the 7075 is about 20% greater than that of the 7050
At 4V of OV where the T.E. is about 97% at the peak we found a “never achieved” PDE of about 78%!!! Compatible with a FF of 84% and a QE of 95%
54%
74%
78%

## Slide 12
PDE@405nm vs OCT
G. Bonanno – INAF O.A. Catania – LIGTH 17 – Ringberg Castle, 19 – October 2017
PDE versus OCT is evaluated at two temperatures 2°C and 15°C
15°C is the ASTRI camera working temperature
A FEE able to operate a good PZC will allow lower OCT independently from temperature
At 2°C the DCR is lowered and the pile-up effect is reduced

## Slide 13
Relative PDE of the LVR3 respect to the LVR2
G. Bonanno – INAF O.A. Catania – LIGTH 17 – Ringberg Castle, 19 – October 2017

## Slide 14
Effect of the IR filter on the OCT
G. Bonanno – INAF O.A. Catania – LIGTH 17 – Ringberg Castle, 19 – October 2017
Three plates IR-Filter
Minimum distance in mm between the central SiPMs and the first filter plate

## Slide 15
Conclusions
G. Bonanno – INAF O.A. Catania – LIGTH 17 – Ringberg Castle, 19 – October 2017
The characterization carried out demonstrates that the 7x7 mm2 with 75µm microcell SiPM can achieve very high PDE and low OCT
The OCT essentially depends on:
Sensitive area dimensions
Coating on top of the sensitive area
Optics in front of the detector
In the ASTRI camera we have to take into account two more effects :
Pile-up due to the operating temperature or better to say to the DCR
Inefficient Pole-Zero Cancellation (PZC) of the Front End Electronics (FEE)
A FEE with an appropriate PZC will remove these two effects.
At 3V of overvoltage a PDE of 74% at a peak has been found, and even better at 4V, where the OCT is less than 10%, an impressive peak of about 78% has been measured.
A new LVR3 series has been manufactured that improves the response in the 285 – 380 nm spectral range.