## Slide 1: Imaging capabilities of the EDET-80k detector
Ibrahym Dourki
22nd International Workshop on DEPFET Detectors and Applications
8-1 April 2018, Ringberg Castle, Germany
Ibrahym Dourki, PhD Student
Max Planck Institute for the Structure and Dynamics of Matter, Center for Free Electron Laser Science (CFEL) Luruper Chausee 149 D-22761 Hamburg (Germany)
ibrahym.dki@gmail.com
ibrahym.dourki@mpsd.mpg.de

## Slide 2: Outline
Introduction: TEM, Bio samples and radiation
Simulation of detector performance
Charge sharing
Signal processing
Noise
Imaging capability of EDET: DQE
Conclusion and outlook

## Slide 3
Keskin et al. J. Phys. Chem. Lett. 6, 4487−4492 (2015)
Introduction
DNA self assembly
200 nm
Goal: Real space and real time imaging using TEM.
Fast molecular dynamics in biology:
Need direct and fast detector: EDET-80kHz
80 000 Hz
Some phenomena classified by spatial and temporal resolutions [1]
W. E. King et al., J. Appl. Phys. 97, 111101, (2005).

## Slide 4
Low dose to avoid radiation damage of the sample
- Need direct detector with large DQE
Large charge handling capability:
- Thin sensor (e.g. 50 µm) with 100 % Detection efficiency
Single electron detection with high probability
Fast Readout: 80 kHz
- 100 frames form one movie of 1.28 ms
I. Dourki et al. 2017 JINST 12 C03047
512x512 pixels
Shielding
e-

## Slide 5
Tools:
Monte Carlo simulation: detector modeling
DEPFET technology concept with signal compression
fulfill these requirements
M. Porro et al. , IEEE Trans. Nucl Sci. (2012)
P. Lechner et al. , IEEE Nuc. Sci. Sym. (2011)
DEPFET response and the number of ADC bits are the key factors determining the achievable dynamic range of EDET-80
Fast saturation and limited DR
Single e- resolution and large DR
gq=2.1064x10-4 µA/e-
gq=3.7133x10-5 µA/e-
Rainer, HLL

## Slide 6: TEM movie
Dynamic range limit: Speed is the goal….not beauty
Lit.: G. McMullan and a. R. Faruqi, Nucl. Instruments Methods Phys. Res. Sect. A Accel. Spectrometers, Detect. Assoc. Equip. 591, 129 (2008).
0.5 mu
0.4 e-/pixel (scaled)
40 e-/pixel (scaled)
3500 e-/pixel (scaled)
16000 e-/pixel (scaled)
►100 el/px Poisson only
►100 el/px Poisson only

## Slide 7
512 x 512
512 x 512
512 x 512
512 x 512
Detector specification and simulation
Direct hit electron detector
Megapixel detector made up of four quadrants (512x512)
Pixel size: 60 µm x 60 µm
Full frame recording at 80kHz (Four-fold rolling shutter mode)
High statistical probability to discriminate single electron hits
High detection efficiency
Applications: Fast Imaging Experiments (e.g. TEM)
Dynamic range > 200 primaries
Switcher
DCD
DMC
512x512 pixels
Shielding
e-
Digitization: 8 bit resolution
Notes: 1) We changed: quantum -> detection
Switcher:
1- Switch on / off the DEPFET
2- Switcher performs also clear operation of the DEPFET via switching between two voltages  clear pulse
DCD: Drain Current Digitizer: Digitize the drain current from matrix
DHP (Data Handeling Processor) : Data Handeling of the digital domain. It is responsible for data reduction (reduces the amount of data via zero suppresion of non-hit pixels and assigns the time stamps to the event) and send data off the module to the backend electronics.

## Slide 8: EDET DATA
Fundamental statistics of process
e-h pair creation
e-h pair creation
Source
Poisson noise (G4)
Landau noise (G4)
Fano noise (ext.)
Obsevable
Digitization
Look-up tables
Backscatter (G4)
Charge coll. (ext.)
Analog noise DCD (est.)
ADU
*b. scatter correction
?
ADU | e- (primary) | +err | -err
0 | 0 | 1 | 0.5
1 | 0.2 | 0.4 | 0.1
... |  |  | 
255 | 150 | 19 | 18
DCDE gain k1
Granularity
ESTIMATOR
Biased Analysis
DAQ
Calibration
DepFET
Current noise (est.)
Biased Results
*thermal & leak
*read out
*cosmic (-)
!
*pedestals
Experiment
Simulation

## Slide 9
35 mm gap
Sloped wall
Straight wall
Backscattering reduction
Detector design optimization
Simulation of detector performance

## Slide 10
55.7 %
Charge sharing

## Slide 11
1) e-h pairs distribution
2) DEPFET current distribuion
Without Poisson noise
4) Mean of PE in vs. ADU
3) Digitization: ADU distribuion

## Slide 12
Assumption: Electrons hitting detector obey Poisson statistic
Noise: standard deviation:
k: events (k= 1,2,…,N)
m: Average of the number of events
For each Poisson event, a corresponding number of e-h pairs
is calculated: noise is now added to e-h pairs generated in Edet
For 1 event: Sin1 = Si1+ni1 (n1 is the noise)
For 2 events: Sin2 = Si2+ni2
For 3 events: Sin3 = Si3+ni3
. . . .
. . . .
For N events: SinN = SiN+niN
With Poisson noise
e-

## Slide 13
Poisson noise is dominant!
Linear relatioship between PE input and PE out for k<0.667 (full 8bits DR is filled)
Response departs from linearity: k>0.667
Higher gain: single electron resolution
Lower gain: Whole high dynamic range is covered
Poisson noise

## Slide 14
Fano Factor
Fluctuation in e-h pair production
Nr. e-h pairs generated in detector is given by:
for Si
is divided into two fractions.
fraction goes to the ionization: e-h pairs
fraction goes to lattice excitation: photon generation
Fano Factor:
and
Fluctuation around the average number of e-h pairs does not fellow Poisson statistics . It is smaller by a factor of .
e-
e-
+
+
+
e-
e-
photon
+
e-

## Slide 15
e-
dEp1
dEp2
dEp3
dEp4
dEp5
dEp6
Mean = nehs
Draw one random number:
Do the same for other dEpi
Difference calculation for each event (PE):
dEpi : Energy depsoition per step
nehs : Number of e-h pairs generated in this setp
Neh : Total number of e-h pairs
PD(mean, )

## Slide 16
Detector
SNRin
SNRout
Imaging capability of EDET: DQE and MTF
 1
Radiation senistive samples: DQE is a key
2xDQE achieves same SNR with half the dose.
Ideal detector: ⇒ DQE=1
Noise components:
Shot noise (Poisson statistic):
Landau distribution: electrons are detected with different weights
Readout noise, Fano, background, etc.
33 %
M. Kuijper et al. / Journal of Structural Biology 192 (2015) 179–187

## Slide 17: Quantization of DQE
Infinit pixel
Cluster
Single hits
DQE veruse dose

## Slide 18
DQE(f) in frequency domain can be calculated from the following equation:
MTF measures how well a detector is able to transfer
signals of various SF from the input to the output.
MTFin : Modulation Transfer function at the input
MTFout : Modulation transfer function at the output
MTF is calculated from slanted edge
e-
DQE(0)
  Non-spatial counting capabilities: Infinity-pixel

## Slide 19
DQE(0)
  Non-spatial counting capabilities: Infinity-pixel
G. McMullan et al. / Ultramicroscopy 147 (2014) 156–163
Detector | #pixel | Pixel length [mu] | A per pixel at mag. 50k | Object length at mag. 50k [mu] | Ava. Speed [fps] | Max. Speed [fps] | Sweet flux [e-/px/s]
Falcon III | 20 M | 14 | 2.8 | 0.41 | 40 | -> | 160
K2 Summit | 14 M | 5 | 1 | 0.38 | 400 | -> | <2
DE-20 | 17 M | 6.4 | 1.3 | 0.67 | 25 | -> | 100
EDET_80k | 1 M | 60 | 12 | 1.2 | 1000 | 80 000 | 20 000

## Slide 20
Summary & Outlook
Detailed detector simulation studies have been performed.
A competitive direct detector with improved imaging performance has been identified.
Next steps:
DQE in spatial frequency domain
DQE from single hits
DQE vs. Dose
Dominant source of noise in the detector: Poisson noise
Notes: Summary & Outlook

## Slide 21
HLL ► Ladislaw Andricek ● Martin Hensel ● Christian Koffmane ● Jelena Ninkovic ● Gerhard Schaller ● Martina Schnecke ● Florian Schopper ● Andreas Wassatsch ● Christian Zirr
KIT ► Ivan Peric
USI ► Klaus Gärtner
MPSD ► Ibrahym Dourki ● Sascha Epp ● Djordje Gitaric ● R. J. Dwayne Miller ● Fabian Westermeier
EDET Team: Acknowlegements

## Slide 22
Thank you for your attention