## Slide 1
EDET Hybrid Measurement Results

## Slide 2: OVERVIEW
FROM HYBRID5 to EDET-HYBRID
  PHILOSOPHY
  CHANGES
  COMPLETED COMPONENTS
MAPPING ISSUES
STANDARD MEASUREMENTS
CALIBRATION ISSUES

## Slide 3: Motivation for new E1 Hybrid Board
Space Constraints
  EDET small matrix (active area: 12x7mm2) larger than BELLE small matrix (7.68x3.84mm2) permitting connections to all DCDE and Switcher channels  new wirebond adapter
Integrate new connections
  compatibility to BELLE PS & PSP  Glenair connectors
  Switcher substrate line, Switcher Vref line
  enabling sense lines to all ASIC voltages
Decision on data interface
  Change 2nd infiniband (JTAG lines) connector to RJ45 in order to get rid of JTAG breakout board

## Slide 4: Philosophy
Starting with the functional BELLE II chain all components are stepwise replaced and tested
Hybrid5 Board
DHP
DCD
SW
B-matrix
BELLE II
PS
JTAG
Break-
Out
Board
DHE
DHE
Samtec
Infini
band
Infini
band
Infini
band
RJ45
RJ45
Infini
band
Infini
band
Infini
band

## Slide 5: CHANGE (I) – DCDE
Starting with the functional BELLE II chain all components are stepwise replaced and tested
Hybrid5 Board
DHP
DCDE
SW
B-matrix
BELLE II
PS
JTAG
Break-
Out
Board
DHE
DHE
Samtec
Infini
band
Infini
band
Infini
band
RJ45
RJ45
Infini
band
Infini
band
Infini
band

## Slide 6: CHANGE (II) – EDET MATRIX
Starting with the functional BELLE II chain all components are stepwise replaced and tested
Hybrid5 Board
DHP
DCDE
SW
E-matrix
BELLE II
PS
JTAG
Break-
Out
Board
DHE
DHE
Samtec
Infini
band
Infini
band
Infini
band
RJ45
RJ45
Infini
band
Infini
band
Infini
band

## Slide 7: CHANGE (III) – EDET HYBRID BOARD
Starting with the functional BELLE II chain all components are stepwise replaced and tested
E1 Hybrid Board
DHP
DCDE
SW
E-matrix
BELLE II
PS
JTAG
Break-
Out
Board
DHE
DHE
Samtec
RJ45
RJ45
Infini
band
Infini
band

## Slide 8: CHANGE (IV) – PSP
Starting with the functional BELLE II chain all components are stepwise replaced and tested
E1 Hybrid Board
DHP
DCDE
SW
E-matrix
PSP
JTAG
Break-
Out
Board
DHE
DHE
Glenair Micro D 51
RJ45
RJ45
Infini
band
Infini
band
Glenair Micro D 51
PSP
Breakout-Board

## Slide 9: E1 – I – 01

## Slide 10: PSP Breakout-Board – Final Board

## Slide 11: Wire-Bonding then (BELLE II) …
DOWN
UP
DOWN
UP
UP
DOWN
UP
DOWN

## Slide 12: … and today (small EDET matrix) same WBA
UP
DOWN
UP
DOWN
UP
DOWN
UP
DOWN

## Slide 13: MAPPING APPROACH (1)
wire-bonding starts with closest pad rows, i.e. 1->1, 2->2, 3->3, 4->4
0 1 2 3
3 2 1 0

## Slide 14: MAPPING APPROACH (2)
enter correct pad connections from the layout in an excel sheet
electrical row consists of 4 geometrical rows
the pad arrangement for the EDET-Small-Matrix does not fully correspond to the geometrical numbering: 3,2,1,0  3,1,2,0
sort by DCD channel index & reversely by matrix index and copy corresponding matrix to python program mapping.py
Matrix Arrangement |  |  | 
255 | 253 | 254 | 252
251 | 249 | 250 | 248
247 | 245 | 246 | 244
243 | 241 | 242 | 240
239 | 237 | 238 | 236
235 | 233 | 234 | 232
231 | 229 | 230 | 228
227 | 225 | 226 | 224
223 | 221 | 222 | 220
219 | 217 | 218 | 216
215 | 213 | 214 | 212
211 | 209 | 210 | 208
207 | 205 | 206 | 204
203 | 201 | 202 | 200
199 | 197 | 198 | 196
195 | 193 | 194 | 192
191 | 189 | 190 | 188
187 | 185 | 186 | 184
183 | 181 | 182 | 180
179 | 177 | 178 | 176
175 | 173 | 174 | 172
171 | 169 | 170 | 168
167 | 165 | 166 | 164
163 | 161 | 162 | 160
159 | 157 | 158 | 156
155 | 153 | 154 | 152
151 | 149 | 150 | 148
147 | 145 | 146 | 144
143 | 141 | 142 | 140
139 | 137 | 138 | 136
135 | 133 | 134 | 132
131 | 129 | 130 | 128
127 | 125 | 126 | 124
123 | 121 | 122 | 120
119 | 117 | 118 | 116
115 | 113 | 114 | 112
111 | 109 | 110 | 108
107 | 105 | 106 | 104
103 | 101 | 102 | 100
99 | 97 | 98 | 96
95 | 93 | 94 | 92
91 | 89 | 90 | 88
87 | 85 | 86 | 84
83 | 81 | 82 | 80
79 | 77 | 78 | 76
75 | 73 | 74 | 72
71 | 69 | 70 | 68
67 | 65 | 66 | 64
63 | 61 | 62 | 60
59 | 57 | 58 | 56
55 | 53 | 54 | 52
51 | 49 | 50 | 48
47 | 45 | 46 | 44
43 | 41 | 42 | 40
39 | 37 | 38 | 36
35 | 33 | 34 | 32
31 | 29 | 30 | 28
27 | 25 | 26 | 24
23 | 21 | 22 | 20
19 | 17 | 18 | 16
15 | 13 | 14 | 12
11 | 9 | 10 | 8
7 | 5 | 6 | 4
3 | 1 | 2 | 0
DCD WB |  |  | 
0 |  | 16 | 
64 | 32 | 80 | 48
1 | 96 | 17 | 112
65 | 33 | 81 | 49
2 | 97 | 18 | 113
66 | 34 | 82 | 50
3 | 98 | 19 | 114
67 | 35 | 83 | 51
4 | 99 | 20 | 115
68 | 36 | 84 | 52
5 | 100 | 21 | 116
69 | 37 | 85 | 53
6 | 101 | 22 | 117
70 | 38 | 86 | 54
15 | 102 | 13 | 118
7 | 11 | 8 | 9
14 | 10 | 31 | 12
25 | 29 | 23 | 27
28 | 24 | 30 | 26
43 | 47 | 41 | 45
42 | 39 | 44 | 40
51 | 46 | 59 | 63
56 | 57 | 58 | 55
79 | 60 | 77 | 62
71 | 75 | 72 | 73
78 | 74 | 95 | 76
89 | 93 | 87 | 91
92 | 88 | 94 | 90
107 | 111 | 105 | 109
106 | 103 | 108 | 104
125 | 110 | 123 | 127
120 | 121 | 122 | 119
143 | 124 | 141 | 126
135 | 139 | 136 | 137
142 | 138 | 159 | 140
153 | 157 | 151 | 155
156 | 152 | 158 | 154
171 | 175 | 169 | 173
170 | 167 | 172 | 168
189 | 174 | 187 | 191
184 | 185 | 186 | 183
207 | 188 | 205 | 190
199 | 203 | 200 | 201
206 | 202 | 223 | 204
217 | 221 | 215 | 219
220 | 216 | 222 | 218
235 | 239 | 233 | 237
234 | 231 | 236 | 232
253 | 238 | 251 | 255
248 | 249 | 250 | 247
134 | 252 | 150 | 254
198 | 166 | 214 | 182
133 | 230 | 149 | 246
197 | 165 | 213 | 181
132 | 229 | 148 | 245
196 | 164 | 212 | 180
131 | 228 | 147 | 244
195 | 163 | 211 | 179
130 | 227 | 146 | 243
194 | 162 | 210 | 178
129 | 226 | 145 | 242
193 | 161 | 209 | 177
128 | 225 | 144 | 241
192 | 160 | 208 | 176
 | 224 |  | 240

## Slide 15: MAPPING – FUNCTIONAL TEST
E1-I-02 Hybrid Board

## Slide 16: Laser Spot x-direction

## Slide 17: DCDE-Manual DAC-Registers

## Slide 18: DCDE Gain Variation
En30 | En60 | En90 | En120 | Gain | Times lowest Gain
1 | 1 | 1 | 1 | 0.061 | 1.0
0 | 1 | 1 | 1 | 0.063 | 1.0
1 | 1 | 1 | 0 | 0.065 | 1.1
0 | 1 | 1 | 0 | 0.067 | 1.1
1 | 0 | 1 | 1 | 0.087 | 1.4
0 | 0 | 1 | 1 | 0.091 | 1.5
1 | 0 | 1 | 0 | 0.095 | 1.6
0 | 0 | 1 | 0 | 0.100 | 1.7
1 | 1 | 0 | 1 | 0.154 | 2.5
0 | 1 | 0 | 1 | 0.167 | 2.8
1 | 1 | 0 | 0 | 0.182 | 3.0
0 | 1 | 0 | 0 | 0.200 | 3.3
1 | 0 | 0 | 1 | 0.667 | 11.0
0 | 0 | 0 | 1 | 1.000 | 16.5
1 | 0 | 0 | 0 | 2 | 33
Rf: En30 = 30k
En60 = 3k
En90 = 1.5k
En120 = 15k
Rs = 15k
Gain | dynamic range for 200 ADU [µA]
1 | 380
1.5 | 250
3 | 137
11 | 41
33 | 15
Resistor | EDET | BELLE
En30 | 30k | 19k
En60 | 3k | 15k
En90 | 1.5k | 26k
En120 | 15k | not used

## Slide 19: BELLE (top) vs EDET (bottom) PS – low gain
DHE current source limited to 75µA  solution: compound curves with IPDAC

## Slide 20: BELLE (top) vs EDET (bottom) PS – high gain

## Slide 21: Matrix based Transfer Curves – high gain

## Slide 22: DELAY SCAN – E1-I-04

## Slide 23: CALIBRATION BRAINSTORMING
METHOD
DESCRIPTION
PROS
CONS
X
-
ray Source
(low energy)
Cd
109
22keV
•
acceptable interaction
probability
•
equipment
HLL
available
•
generates only
about
6000
•
geometry allows for only 1
-
2
events per readout
X
-
ray Source
(high energy)
beyond 140keV
(corresponds to 10ADUs)
•
no vacuum
apparatus necessary
•
very low photoabsorption
probability of 0.01%
•
Compton effect dominates
•
availability of Source

-
emitter
Sr
90
continuous spectrum
•
generates over 100,000 e/h pairs
(I. Dourki)
•
around
300 keV
electrons would leave the
detector, so escape energy has
to be measured (scintillator)

emitter
Am
241
5.486MeV
(encapsulated as 60keV
g
-
source usable)
•
forces
internal gate capacity to
maximum level (1,000,000
)
•
vacuum
device and
corresponding interconnection
hardware necessary
•
detector radiation damages
•
radiation protection
N
SEM/TEM
•
gradual increase of both
intensity (in 10eV steps up to
300keV) and quantity of
•
vacuum
device and
corresponding interconnection
hardware necessary
Laser
660nm
•
easy focusing
•
intensity
controllable via applied
voltage and time (pulsing)
•
equipment HLL available
•
exact energy transfer unknown
(approximation with optical
power meter)
•
only cell
-
by
-
cell calibration
CLEAR Backinjection
capitalizes on
variation of
integration time
•
relatively
quick calibration of all
pixels possible
•
level of original backinjection not
known (but can be measured
indirectly
–
laser)
•
finding
optimal operation point
for all pixels (process variation)

## Slide 24: Above 60keV Compton Scattering dominates
Chart: Photon Energy [keV]
| 1 | 1.5 | 1.839 | 1.839 | 2 | 3 | 4 | 5 | 6 | 8 | 10 | 15 | 20 | 30 | 40 | 50 | 60 | 80 | 100
Si Mass Attenuation | 1570 | 535.6 | 309.2 | 3193 | 2777 | 978.4 | 452.8 | 245.1 | 147 | 64.69 | 33.88 | 10.34 | 4.463 | 1.436 | 0.701 | 0.4385 | 0.3206 | 0.2228 | 0.1835
Photoelectric Effect | 1567 | 533.3 | 307.1 | 3191 | 2775 | 976.7 | 451.4 | 243.8 | 145.8 | 63.79 | 33.15 | 9.848 | 4.089 | 1.161 | 0.4687 | 0.2307 | 0.1288 | 0.0512 | 0.02498
Compton Scattering | 0.01317 | 0.02393 | 0.03081 | 0.03081 | 0.03388 | 0.04962 | 0.06135 | 0.0711 | 0.07983 | 0.09507 | 0.1076 | 0.1289 | 0.1402 | 0.1501 | 0.1534 | 0.1538 | 0.1526 | 0.1483 | 0.1432
Source: NIST data

## Slide 25: Different Energetic Primary Electrons
Ibrahym Dourki

## Slide 26: Spectra of high energetic electrons leaving Si
Ibrahym Dourki

## Slide 27: average path length of electrons in Si
Ibrahym Dourki

## Slide 28: SEM/TEM – JEOL

## Slide 29: SOLUTION
Calibrate laser pulse with Cd109 source in high gain mode (from BELLE experience about 20 ADU, which corresponds to 1-2 ADU in lowest gain mode)
Step-increase of pulse count within one readout period in order to measure a discrete gain curve

## Slide 30: Find optimal HV-Voltage 0V,-10V, -20V, -30V

## Slide 31: Expected Cd109-Peak – Calculations
parameter | value | unit
L(small matrix gate) | 4.4* | µm
L(small pixel gate) | 5.4* | µm
W(small matrix gate) | 28.6* | µm
W(small pixel gate) | 21.6* | µm
DCD-E dynamic range (high gain) | 20 | µA
 1 ADU | 78.4 | nA
Cd109 – 22 keV | 6000 | e-
small pixel measured gq | 200 | pA/e-
6000 x 200pA  1.2µA | 15.3 | ADU
, | 74, 100(140) | mA
 rescaled by and gq | 15.3 (13.0) | ADU
parameter | value | unit
L(small matrix gate) | 4.4* | µm
L(small pixel gate) | 5.4* | µm
W(small matrix gate) | 28.6* | µm
W(small pixel gate) | 21.6* | µm
DCD-E dynamic range (high gain) | 20 | µA
 1 ADU | 78.4 | nA
Cd109 – 22 keV | 6000 | e-
small pixel measured gq | 200 | pA/e-
6000 x 200pA  1.2µA | 15.3 | ADU
 | 74, 100(140) | mA
 | 15.3 (13.0) | ADU
* For wet etching normally a value of 1.4µm has to be deducted (in case of L) or added (in case of W) respectively from the design value; here it was plasma-etched  only 0.6µm

## Slide 32: Calculated hits with Cd-Source

## Slide 33: Find optimal Voltages: high offset dispersion
GateOn= -1,660mV
ClearOn=18,000mV
ClearOff=3,000mV
HV= -25,000
higher (vs BELLE) offset dispersion due to
higher gm (32%)
higher gain (15%)
different threshold level
…
E1-I-04 Hybrid Board

## Slide 34: Cd-Source Measurement (1)

## Slide 35: Cd-Source Measurement (2)
CONCLUSION: channel-length of 5µA should be chosen

## Slide 36: NEXT STEPS
Optimize Settings for Source Measurements
Calibration of low gain/signal compression with Laser Pulse
Integration of PSP

## Slide 37
THANK YOU FOR YOUR ATTENTION

## Slide 38: PSP Breakout-Board – electronic schematics
Connector x2
Power Group x7 +
2 x 6-pole jumpers for AGND & DGND
Sense Group x19
Normal Load Group x18
circuit diagram/layout documents the electrical wiring of different components
define footprints and functionalities in component library
PCB standard size (W x L x H)
100mm x 160mm x 1.6mm

## Slide 39: PSP Breakout-Board - Design
FC/PC Terminated Fiber

## Slide 40: DCDE Noise
1 primary electron generates
in 50µm thick silicon detector
EDET detector response for
Incremental Signal for
Gain | range 200 ADU [µA] | Current/ADU [µA] | DCDE noise [ADU] | DCDE noise [µA]
1 | 380 | 1.9 | 0.92* | 1.75
1.5 | 250 | 1.25 | 0.69 | 0.86
3 | 137 | 0.69 | 0.69 | 0.48
11 | 41 | 0.2 | 0.96 | 0.19
33 | 15 | 0.08 | 1.50 | 0.12
*Maybe too high because of discrete IPDAC transition
DCDE noise mainly induced
by quantization error (0.5 ADU)
Not relevant because of statistical deviations
for spatiotemporal signal generation
100 | 100 | 100 | 100
100 | 110 | 100 | 100
100 | 110 | 90 | 100
100 | 90 | 100 | 100
Ibrahym Dourki (CFEL)

## Slide 41: SAMPLING POINT SCAN – E1-I-04