## Slide 1: Improving Scanning X-Ray Microscopy
Markus Weigand
Using Silicon Sensors

## Slide 2: Outline
STXM: An overview of an established method
Improving Detectors:
Si APDs for time resolved imaging
Fast in Vacuum CCDs for ptychography
Outlook: Detector Wishlist

## Slide 3: Why X-ray Microscopy?
Resolution set by Abbes Limit:
Basic Idea:
X-rays have 1000 times lower
-> 1000 times better resolution!
Reality:
(soft) X-rays impossible to focus conventionally, requires diffractive optics
Consequences:
Need for Monochromatic, coherent x-ray light requires synchrotron sources

## Slide 4: Scanning Transmission X-ray Microscopy
Probe X-ray absorption on tiny focal spot
Image creation by sequential position scanning
Resolution given by spot size (<25nm)
X-ray absorption techniques as contrast mechanism
Operation Scheme:

## Slide 5: STXM in Real Life

## Slide 6: Photon Detection
P43 Phosphor
Lightguide
Photomultiplier Tube
Saturation at about 15M counts/s
Indirect detection -> Photon energy dependent count rate and noise statistics
Amplification of higher harmonic light pollution
Phosphor decay time O(1ms) -> prevents fast scanning
Very sensitive to magnetic fields
High voltage requires encapsulation
Susceptible to helium poisoning
But: Works Ok with good efficiency even at low photon energies

## Slide 7: Why Soft X-rays?
Photon Energy not relevant for resolution
Low Photon energy -> high optical density
  Thin samples needed
  High elemental contrast for small objects
150-2000 eV can (almost) cover whole periodic table with single endstation configuration
K
L Edge
M

## Slide 8: Analyzing Aerosols
Na pre-edge
Na post-edge
Na map
C map
N map
Ca map
O map
Fe map
Pöhlker et al.,
Science 337 (2012) 1075-1078

## Slide 9: Imaging Battery Materials
      Jongwoo Lim, W. Chueh, Y. Li
      William Chueh Group
      Materials Science and Engineering
Stanford University
SiN
LiFePO4
Transmission
LiFe2+PO4 (Discharged)
Fe3+PO4 (Charged)
1 µm

## Slide 10: Magnetic Imaging
Thickness | Fe L-3 | Co L-3 | Ni L-3 | Gd M-5
0.1 nm | 0.3% | 0.3% | 0.1% | 0.6%
1 nm | 3.3% | 3.1% | 1.2% | 6.4%
5 nm | 17.6% | 16.5% | 6.2% | 36.3%
25 nm | 50.0% | 46.3% | 16.1% | 117.1%
Imaging local magnetization
Element specific imaging
Extremely fast
Very strong contrast:
Notes: Photon helicity erwähnen . > erleutern extraction static contrast
Evtl XMCD bild für domänen

## Slide 11: Time Marches ON
Dramatic increase
in light source brilliance
Increasing demands
from Experiments
Better Resolution
Faster Scanning
Better Image Fidelity
Time Resolved Imaging
Exploiting Beam Coherence
PMTs insufficient
for any of these requirements!

## Slide 12: Si Avalanche Photodiodes
Short recovery times (<1ns), important due to time structure of source
1.5 OOM more linear photon handling capabilities (compared to PMT)
Internal gain: ca 18.5 e- per eV for X-rays
Direct Detection (no phosphor delay) -> Photon tracking capability
SNR >106 @ >108 events/s
Designed for infrared, not x-rays!
Still work when decapped, though!
1300 eV Photons on Hamamatsu S2381
amplified by Femto HSA-Y-2-60

## Slide 13: Energy Dependence of APD
Gain drops dramatic below 600 eV!
QE goes from apprx. 50% at 600ev to <1% at 300eV
-> Unusable for biological samples / ps chemistry!

## Slide 14: Looking at APDs on the nanoscale
Efficiency plot of the surface of an APD: Lower photon energies make hot and cold spots much more pronounced.
At 400eV, factor 10 split of QE!
Shift in peak indicated only 2 photon
events strong enough to be detected
Side Note: remarkable resilience to x-ray exposure:
No sign of degradation even after extended exposure
to photon flux

## Slide 15: Time Resolved Microscopy
e
X
Repetition Rate O(500 MHz)
X-ray Pulses:
Length: <50ps
Insertion Device

## Slide 16: Pump-and-Probe Setup
Time
2ns
System
Response
499.65 MHz
Sync
499.65 MHz
Sync
Photon Detection
Photon Sorting
DDS Pump:
Direct RF or Pulse Trigger
Synchroton
With the right tricks
Free choice of excitation parameters
Simultaneous acquisition of all time channels!
All buckets contribute to all time channels

## Slide 17: Controlling Domain Wall Motion
Purely shot-noise limited,
despite:
Usage of full Filling Pattern
Top-up Operation
Femto-Slicing Bunches
A.Bisig, et. al.
Nature Commun 4, 2328
J-S. Kim, et. al.
Nature Commun 5, 3429
601 Channels
1.2us Period
Raw Data
20 minute
measurement

## Slide 18: Large-Scale Dynamic Imaging
Multi GHz Spinwaves
in Magnonic Crystals
J. Gräfe, MPI-IS
23 Ch, 300ps Period
16*8 um, 20nm/pixel
Phase
Amplitude
Notes: Erwähnen 16*8um, 20nm pixel, 5h.

## Slide 19: Large-Scale Dynamic Imaging
Ga:YIG Ferrimagnet Sublattice Imaging
S. Mayr et al. (in preparation).
25 µm
Spin Waves | 3.5 GHz | 170 mT
Normalized
Domain Walls | 118 MHz | 100 mT
Diagonal stripes due to substrate thickness variations, can be normalized out
Orders of magnitude of frequency range can be covered in single setup
Unique way to directly image dynamic magnetic behavior!

## Slide 20: Improving Resolution
The whole point of microscopy, is you think about it…
Theoretical Limit
O(1nm) for soft X-rays
Practical Limits:
Zone Plate Production
Mechanical Stability
Focal Length requirements
(Radiation Dosage)
Reality: 15-20 nm resolution
Not going to happen

## Slide 21: Recording Sample Transmission
Intensity behind Sample
Sample Diffraction can be used to improve Resolution!
CCD
Keyword: Ptychography
10.1126/science.1158573
Chart: Publication Hits
| 2006 | 2007 | 2008 | 2009 | 2010 | 2011 | 2012 | 2013 | 2014 | 2015
Web of Knowledge | 0.01 | 1 | 6 | 6 | 18 | 27 | 26 | 55 | 76 | 94

## Slide 22: Camera Requirements
Soft X-rays == In vaccum camera
High dynamic range
Fast readout speed (ideally >100 Hz)
Large photon handling capability (>10^9 ph/s @ 1keV)
Large chip (for covering big numeric apertures)
High resolution is NOT a priority
PnSensor pnCCD
450 fps
264 x 264 pixels (48um size)
Single Photon Detection
About 12 bit dynamic range

## Slide 23: Dynamic stacking
Averaged signal
Frames 450 Hz
Pixels
4-5 bit of additional dynamic range for 100ms dwell
Pixel delay 10 ms
Full-well capacity too low for good dynamic range even at lowest gain range -> Need for stacking

## Slide 24: Making Output Useable
Raw output 132*528 pixel
CAMEX 1
CAMEX 2
CAMEX 3
CAMEX 4
4 readout chips
132 parallel ADC each

## Slide 25: Making Output Useable II
No Global Shutter
Frame transfer from center to both sides
Individual lines have different offset
Can be reduced pretty well using weighted median subtraction, but still situationally problematic

## Slide 26: Improving Spatial Resolution (using „sharp camera“)
Classical STXM
Ptychography
100 nm
Real Beam
Pattern
Flat Contrast Transfer Function
Notes: Fehlt die Take-home message

## Slide 27: Ptychography Tests at MAXYMUS
      Jongwoo Lim, W. Chueh, Y. Li
      William Chueh Group
      Materials Science and Engineering
Stanford University
Fe2+
Fe3+
500 nm
1 um
200 nm
Ca 6 nm half-pitch Resolution with 100nm FZP
Illuminating spot size
Benefits from rapid development of X-ray cameras for FELs and upcoming high brilliance sources
Towards few nm resolution
chemical microscopy!

## Slide 28: For the future
APD viable for time resolved detection, but could be improved:
Do not apply SiO2 AR layer
Use optically opaque, but x-ray transparent top contact (i.e. 10nm Al, etc.)
Optimize thickness of avalanche layer
Better control of doping homogeneity
Next Generation CCDs
In development with HLL
1024^2 pixels on a 36.8mm^2 chip
Compared to existing camera will allow:
  3x resolution
  10x light gathering
  Any combination
Still no global shutter, also no time resolution
The “perfect Detector”:
Combining ptychography with pump and probe
Large sensitive area
Global shutter
Hybrid design for capturing >10^9 events/s with <1ns time resolution

## Slide 29: Thanks to:
  K. Litzius, A. Bisig, M. Kläui
MPI Stuttgart & MAXYMUS
  I. Bykova, J. Gräfe ,E.Goering, M.Bechtel, H. Stoll, K. Keskinbora, U. Sanli, G. Dieterle, J. Förster, C. Stahl, S. Ruoß, J. Simmendinger, M. Kammerer, M. Noske, M. Curcic, L. Bommer, P. Audehm, C. Wolter, B. Baretzky, K. Fauth, G. Schütz
  B. Van Waeyenberge
  S. Wintz, R. Follath, J. Raabe
  D. Shapiro
  S. Raoux, G. Schneider, S. Rehbein, S. Werner, P. Guttmann, R. Müller, M. Mast, S. Hendel, T. Schlegel-Niedrig
  C. Pratsch, T. Lunkenbein
  S. Eisebitt, B. Pfau, I. Will