## Slide 1
EDET – Key Findings
EDET DH80k: Optimization, Calibration & Evaluation
Mai 17th 2022

## Slide 2: SYSTEM OPTIMIZATION What to optimize?
Single Components: DCDE, DEPFET-Sensor, Power System, ...
Interchip Communication: Delay Scans, JTAG
Whole System: Offset Compression, Ghost Charge, Sampling Point Analysis, …
 ULTIMATE GOAL: increase SNR

## Slide 3: SYSTEM OPTIMIZATION Systematic Approach
WORK STEP | Example 1: DCDE | Example 2: Sensor Charge Collection
Define Objective Function | Missing Codes, INL, DNL, noise …  Hierarachy? | Maximize Charge Collection
Define Objective Parameters | Over 15 DCDE 7-bit DAC settings | DEPFET Voltages
Select Methods and suitable Instruments (sometimes have to be developed) | External current source increasing signal injection | 2D-Voltage Sweeps + LED Charge Injection
Program necessary Algorithms |  | 
Output | Transfer Curves | 2-D Charge Collection Maps
Time Management: sweep step widths, selection of parameters, selection of favorable regions, …

## Slide 4: E1_I_2 Can these Transfer Curves Improved?

## Slide 5: E1_I_2 DCDE Optimized

## Slide 6: GHOST CHARGE
Phenomenon of an additional measured pixel current without any external charge injection

## Slide 7: GHOST CHARGE
Can not be calibrated away!

## Slide 8: GHOST CHARGE detailed view

## Slide 9: DIFFERENT MASKS – GHOST CHARGE REMAINS

## Slide 10: ANOTHER EXAMPLE OF GHOST CHARGE

## Slide 11: MITIGATING MEASURES …
Variation Switcher Parameters: extend clear period, increase clear voltage, different sampling points  no impact
Decrease Matrix Current: no impact (corresponding adjustment VnSubIn)
Variate DCDE Parameters: set VTCSFN and VTCCASC to maximum  😁

## Slide 12: GHOST CHARGE after DCDE ADJUSTMENTS

## Slide 13: Pixel Detectors need to be calibrated!
Despite "flat-field" LED-Illumination:
"mosaic" picture instead of homogeneous image

## Slide 14: Cd109 Source Pixel Histograms
For some pixels peak at 25 ADU for others at 45 ADU  need for Calibration

## Slide 15: Cd109 Source Calibration (I)

## Slide 16: Cd109 Source Calibration (II)

## Slide 17: Cd109 Source Calibration (III)

## Slide 18: EDET SIGNAL COMPRESSION EDET Charge Handling Capacity significantly increased

## Slide 19: Calibrate Basis-LED-Intensity with 109Cd Source

## Slide 20: LED allows to establish Charge Curves for all Pixels at once…

## Slide 21: WHOLE SENSOR CALIBRATION

## Slide 22: KNIFE EDGE MEASUREMENT
PHOTOGRAPHY: ISO 12233 (2017)

## Slide 23: DETERMINE EXACT KNIFE EDGE ANGLE

## Slide 24: SUPERSAMPLING & ESTABLISH ESF
Binning

## Slide 25: ESF & LSF

## Slide 26: MODULATION TRANSFER FUNCTION – MTF
In analogy to audio signals where Fourier Transformation switches from time to frequency domain, an image's contrast is decomposed to its spatial frequencies  lp/mm

## Slide 27
THANK YOU FOR YOUR ATTENTION

## Slide 28: EDET HYBRID BOARD
Delay Scan, DCDE optimization, Sample Point Determination, Noise Analysis, Transfer Curves…

## Slide 29: Laser Scan of one single EDET Pixel constant laser charge, 1 µm motorstage steps

## Slide 30: Laser Scan 9 x 16 pixel array constant laser charge
3 µm step
highest pixel value
small tilt
thicker blue=drift line